Standard Operating Procedure for the Processing of EEM’s Samples from Aqualog with R Studio

Prepared by Katie Wampler

Scope and Application

This procedure is used process samples obtained from the Horiba Aqualog. Samples are cleaned and processed, then they’re plotted, indices are calculated, and data is saved.

File Structure

All raw files from the Aqualog should be saved in the T drive (T:\Research\Aqualog_Data) under the appropriate project folder.

Aqualog daily testing runs will be under the validation project folder.

Within each project folder is a folder for each day the Aqualog was run, using the format YYYY_MM_DD (e.g. 2022_08_20).

In the folder for each day if it is:

Unprocessed, it should contain the metadata table, the raw files, and the project folder(s).

Processed, it should contain 5 folders, a zip file, the metadata table, and the project folder(s).

R Basics

When coding in R, everything is case sensitive (e.g. “variable” isn’t the same as “Variable”).

In general, avoid adding spaces between words no matter what, instead of spaces use “_”.

Processing Samples Run Manually

Ensure you have R and R Studio loaded on your computer.

Open the the “EEM's Processing_[most recent date here].R” script (T:\Research\Aqualog_Data\0_Codes and Files\).

NOTE: If you plan to process samples frequently, it’s recommended you save a copy of the code to your personal files in case you make any changes, it won’t affect the main script.

In the script highlight and run all the libraries to load. If this is your first time running the code you may need to install the packages.

Set up the code for your samples

Set the date of the run, this should match the date of the file with the data you want to process. The date should have the form “YYYY-MM-DD”.

Set the project where your data is located (e.g. “2_PNNL_DOM/”)

Specify if the samples ran manually or using the sample Q. If you used the sample Q put “SampleQ”, otherwise put “Manual”. If you are unsure, you can guess, the code will double check for you.

Run lines 28-33, this will set the location of your data.

If you have DOC data, you can run the optional section which will fill in the raw DOC data into your metadata table.

The raw DOC data needs to be saved here: T:/LabMembers/10_Katie/1_Research/2_PNNL_EWEB/Data/Raw DOC Results/, or update the location of your DOC data.

The file name needs to have the form YYYYMMDD_MC_NPOC_TN_data.csv. Where the YYYYMMDD is the day the data was collected in the field.

Update the collectdate variable to the YYYYMMDD from the file, the date the data was collected in the field.

Click the arrow by the section header to collapse the section, then highlight and click run.

Run Section 0. This section contains user functions that allow the code to run. No changes are needed here.

Click the arrow by the section header to collapse.

Highlight and run.

Open Section 1: Code Settings. This section contains user preferences and allows you to modify the code outputs.

The special_notes variable, within the quotations enter any notes you want saved in the processing log. Any errors, deviations from procedures, etc. should go here.

Processing Options

      - The rayleigh_mask variable allows you to specify the width to remove around the rayleigh line. This variable needs to be 4 numbers, the first is the width above the first order rayleigh line, the second is the width below the first order rayleigh line. The third and fourth numbers are the width above and below the second order rayleigh line.

      - The rayleigh_interp variable is two values. It can only be T or F (true or false) which allows you to specify if you should interpolate the missing data for the first order rayleigh line (first value) and second order line (second value).

        - NOTE: Due to the way the interpolation function works and the way the code was written you if you set the first order line to interpolate it will automatically interpolate the second order even if the code specifies F.

        - NOTE: Because the rayleigh lines are usually pretty thick, interpolation often looks weird, so it’s recommended especially for the first order to just remove and leave the values as NA’s.

      - raman_mask and raman_interp are the same as the raleigh_mask and rayleigh_interp variables expect they mask the first and second order raman lines.

        - NOTE: The raman lines are typically much thinner and interpolation is recommended for these lines.

      - The IFE variable can be set to T or F (true or false) and specifies if the samples should be corrected for inner filter effects.

        - Typically, these effects are only important which absorbance values are high (>~1) however, it is recommended to leave these set to T for consistency.

      - The raman_norm variable can be T or F (true or false). This uses the RSU adjust value listed in the metadata table to normalize the data for different integration times. It’s recommended to leave this set to T always.

      - The DOC_norm variable allows the user to normalize the absorbance and EEM’s sample by DOC concentration. This can be set to T or F (true or false). This will only work if the metadata table includes data in the DOC column. This will only affect the data saved as processed files.

Plotting Options

      - The em_clip variable is two numbers specifying the lower and upper limit of the emissions wavelengths to use for plotting.

      - The ex_clip variable is two numbers specifying the lower and upper limit of the emissions wavelengths to use for plotting.

      - The individual variable is a T or F (true or false) which specifies if the code should save the EEM’s plots as individual .png files.

      - The DOC_norm_plot variable is a T or F (true or false) which specifies if the plots should be normalized for DOC concentration.

      - The overall variable is a T or F (true or false) which specifies if the code should save the EEM’s plots as a summary .png file.

      - The overall_ncol variable

      - allows the user to input a number for the number of columns to arrange samples in for the overall summary plot.

      - The overall_title variable has the options of “sampleID” or “description” which specifies what should be used as the title for figures in the summary plot. If “sampleID” it will use the sample ID, if it is set to “description” it will use the description in the metadata table.

      - The nbins variable is a number specifying the number of bins to use for the contour plot of the EEM’s. It’s recommended to use between 5-15, but it should work with any number.

      - The plottitle variable is a T or F (true or false) and allows the user to specify if the individual plots should have a title. If T, it will put on the description from the metadata table.

      - The manualscale variable is a T or F (true or false). If it’s set to T it will use the minimum and maximum values set as manualmax and manualmin. If it’s set to F it will use the minimum and maximum for each individual sample. Using a manual scale works well when you want to compare data and it’s similar in intensity.

Set indices settings

      - The DOC_norm_index variable has the following options: T or F or “both”. If it’s set to T, the indices will be normalized for DOC. If it’s set to F, the indices will not be normalized. If it’s set to “both” it will return both sets. It will return an error when running the code if this is set to T or “both” and there is no DOC data in the metadata table.

      - If the sampsascol variable is set to T, indices will be saved with sample names as columns and index values as rows. If this is set to F it will list samples as rows with indices as columns. The default and recommended value is F.

Optional Skipping of Steps

      - There may be occasions where you want to rerun the code (adding in DOC data, adjusting plotting). To save time, there is options to skip steps to save time when running the code.

      - The raw_to_zip variable is a T or F (true or false). If set to T it will zip any raw .dat files in the main folder into a zip file to preserve the raw data. If set to F it will not zip any files. The recommended setting is F, as the code will override this if there is no existing zip file. You would want to set to T if you’ve run the code, but you’re adding new data to the folder.

      - The calc_indices variable is a T or F (true or false). If T it will recalculate the spectral indices. If F it will not recalculate. This is helpful to set to F when adjusting the plotting to save time when running section 5.

Once all settings are adjusted as desired, click the arrow next to the processing steps section header to collapse and run the section.

Run Sections 2-5 by highlighting and running. This may take a little bit, keep an eye out for error messages.

Once the code has successfully written you will now have processed data in your run folder. The file structure is as follows

1_Absorbance contains the renamed raw absorbance files.

2_Blanks contains the renamed raw blank files the match with the EEM’s files.

3_Samples contains the renames raw EEM sample file.

4_Clean_Absorbance contains the absorbance data saved as .csv files.

5_Processed contains all the cleaned and processed data.

      - Figures contains any plots that were specified in the script

      - There are .csv files with the clean and processed EEM’s data for each sample. These are saved with a unique sample name which is the sampleID_duplicate#_integrationtime

      - OutData_Abs.csv contains the clean and processed absorbance data

      - OutData_FDOM contains the FDOM metrics

      - SpectralIndices contains the absorbance and fluorescence indices calculated

      - Processing_tracking is a text file which lists the processing steps performed with the time and any notes the user added.

At least one .OPJ file, these are the project folder which hold the raw data and are accessible from the Aqualog software.

Metadata table

Zip file with the raw files.

Leave a copy of the data here. All data should remain in the T drive folder. If you would like it elsewhere please copy and paste.

PARAFAC Modeling