Emily Nussdorfer Draft

Speed run (Shortened version for people familiar with process):

Starting

Check if lamp was shut off after night run

-Aqualog software > H2O button > spectra> emission 2D

-Change: wavelength to 550 nm

-Click RTC

-Open shutter slider

-Hold a piece of paper to the left of the cuvette. Green light shows if on

-Set the shutter back to auto. Hit cancel

Turn on Aqualog: warm up about an hour before run

-Record when lamp turned off: Open the most recent aqualog project folder and go to AquaSampleQLog.text and scroll to bottom for time

-Turn on: Flip switch twice if the aqualog was automatically turned off

-Record watershed time in notebook (same as laptop time)

Upload files

-Upload the files saved on the computer into corresponding validation and project folders

Prep

-Get type 1 water: Build resistance to 18.2 MΩ

-Remove samples and tea from the fridge room. Check the date on tea. Get a new one if it's a month past the written date. Buy one from the cafe and get a receipt. If possible, buy extra teas to have on reserve.

Tea standard:

- 24.75 mL of type 1 water

- 250 µL of tea

-Record serial number in notebook

File stuff (easier to do all at once):

-In box navigate to Lab work > Aqualog > instrument-exports

-In both the 1_validation folder and the project folder make a folder named YYYY_MM_DD

-Navigate back to >Aqualog

-On the Lamp Test_YYYY_MM_DD file: click the three dots (more options) > move or copy > instrument-exports > 1_validation > today’s folder. Rename file to today’s date

-Do the same for the Metatable_Manual file: three dots> move or copy > instrument-exports > project > today’s folder. Rename file to today’s date: Metatable_Manual_YYYY_MM_DD

-In the computer files navigate to Aqualog_data 
-In both the 1_validation folder and the project folder make a folder named YYYY_MM_DD

Lamp test:

-Change compartment to basic one without tubes

-Open up aqualog software

-Blue H2O button > spectra > absorbance

Change:

- The data identifier to LampYYMMDD

- Run as a blank only

- Under wavelength settings change high to 700 nm, low to 200 nm and increment to 1nm.

-click run

-Copy lamp test template to today’s date in 1_Validation

-save as Lamp Test_YYYY_MM_DD

-In the Aqualog software, when the project name window pops up, in the validation folder create a new folder for today’s date (YYYY_MM_DD). Name the project as YYYY_MM_DD and hit save

-Copy B(Y) and D(Y) to the lamp test sheet and check if it passes.  If the test shows a fail value, allow the instrument to warm for about another 30 minutes and rerun the test

-Record

Water Raman SNR

-Wipe H20 cuvette with lens wipes

-Insert H2O cuvette into holder

-Software:  collect menu > Aqualog service only menu > Water Raman SNR and Emission Calibration (top)

-data identifier: RSNYYMMDD

-Hit run. Values should say pass. If not wait longer for the Aqualog to heat up

-Record

Raman Scattering Area Unit Tool (RU)

-H2O still in holder

- Software: RU button

Change:

- Data Identifier: RSYYMMDD

- Integration: 10s

- Increment: 2.33 nm (4 pixel)

- CCD Gain: Medium

-Run

-Enter info from Raman Normalize tab; Copy RSU adjust from other tab (DY)

-Rerun again same ID okay

-Check values. Percent difference needs to be below 1%

-Record

Quinine Sulfate Units

-Put H2O cuvette away

-Software: QU button

- change:

- Data Identifier: QSYYMMDD

- Integration (s): 0.05

- Increment (nm): 2.33 nm (4 pixel)

- CCD gain: Medium

- Select sample and blank and collect blank

- Click the three dots next to “collect blank” and save to the validation folder you’ve created for the day using the file name QSBYYYYMMDD.blank, if it doesn’t end in “.blank” it will throw up an error. Hit open

-Run, put QSU blank and QSU sample in when prompted, wiping cuvettes before inserting

-Record

- Copy value in G(Y) column under QSU calculation worksheet tab

-If there is a significant drop (>5000) in the QSU, wipe off the cuvettes with a lens paper

-In the emission sample data tab find the largest value in the K(Y2) column and record the A(X1) value (around 450)

-In the QSU calculation worksheet tab record the value in the A(X) column

Get the Raman Normalization Area

-Replace compartment to the one with tubes

-Software: Sample Q button > select top blank row > open

-Click scan settings, enable injector

-Injector tab > configuration tab > test connection > wash needle > wash loop three times

-Hit cancel and go back to main page

-Software: RU button

Change:

- Data Identifier: RSUYYMMDD (RSU220716)

- Integration: 10s

- Increment: 2.33 nm (4 pixel)

- CCD Gain: Medium

-If doing evening run it will ask for project name- create a new folder for today’s date (YYYY_MM_DDb). Name the project as YYYY_MM_DDb and hit save.

-Record under Raman units in notebooks

-Copy RSU adjust in notebooks and metatable

-Hit save on the toolbar in the Aqualog software to save the project. Then go to file and close. This will close the project with the validation runs.

Run Samples via the Autosampler

-Navigate to the metadata table in the project folder on box

IDs for blanks: BLKYYMMDD#

IDs for teas: postTeaYYMMDD, preTeaYYMMDD

-There is always an instrument blank (not shown on metadata table). If your sample amount is n>9 then you need total 3 blanks, if it’s n = 9-18 it is 4 total blanks

-RSU_area_1s: take the RSU adjust value from the Raman Units page of the Aqualog notebook, paste that here.

-In software > blue Q button > Data/0_Codes and Files/Sample Q Templates)

-Under scan settings ensure the integration time is set to 5s. (1s for USDF_NIFA)

-If you’re running samples in the morning uncheck “Turn Off Xenon lamp After Completion”, if you’re running samples at night ensure this is checked.

- Under samples per blank change to the number of samples you’re running for the day. This should be the maximum index number on your metadata. Repeats per blank are zero. Hit Add.

-Change the Sample ID for the Blank to DIYYMMDD5s (1s for USDF_NIFA)

-Wash needle, wash loop

-Hit save, then save the sample Q file in the project folder under today’s date, name it “MCSN Method_YYYY_MM_DD”, then hit save.

Samples loaded from 1A>1B>1C >2A>2B>2C>……

Pour samples out

-It’s easier to line your samples bottles up in an order that makes sense (eg. chronologically)

- Pre-label your aqualog vials- use a fine point colored sharpie (not black)

-Pour or pipette samples into vials -the volume does not need to be exact- roughly ½ full. (8.5 mL).

-After the samples are put in, triple check the samples are in the same order as the metadata shows

-Place the trays in the autosampler, they should snap in/ lay completely flat. Slide in the cooling tray and close the door

-Before you hit run on the software, ensure the sample chamber with the tubing is in the Aqualog

-Also double check that the lamp is set to turn off if you’re running a batch at night.

-Lastly, check that the output folder is set to today’s date.

-Once this is done, ensure your method is selected and hit “Run”, it will ask you to verify the vial plate, tray, number of samples, and amount of wash water. Please double check these things before hitting OK.

Stay until the blank is collected. As it’s acquiring data, keep an eye on the spectra, it should look like the figure below

Aqualog checklist-. Very rough needs editing

Clean up after last run- morning and evening

- Check if the run completed without errors (look at the sample log).

- Check if the lamp is turned off on the aqualog.

- Upload files from last run

- Take the trays out of the autosampler (only if it slides forward- if it doesn’t then reconnect it to the sampler). Take vials out and put the tray back in (slide cooling cover over it).

- Wipe the caps off with methanol on a kimwipe. Dump vials out in the sink. Rinse both off with DI water. Place caps and vials in “to be acid washed” containers. Note- can do this step at the end while waiting on the first blank to collect.

- If not running any samples turn the autosampler off

- Empty out carboy (below the autosampler) and put output tubing back in

Prep for run- morning and evening

- Check when last rinsed with isopropyl- want to do this regularly ~about each month?

- Remove lipton tea bottle and samples out of the cold room

- A full run is 18 samples

- check the date opened on the tea standard- if over a month get a new one.

- Refill MQ water filter tank and cycle

- Collect type 1 water in it’s labeled bottle

- Refill rinse container  (empty out). Be sure to put tubing in the beaker to avoid contamination.

- Prepare tea standard. Record it in the notebook

- 24.75 mL of type 1 water and 250 uL of tea.

- Type 1 water: Set 5-10 mL pipette to 10.00 and dispense two times; set 0-5ml pipette to 4.750 and dispense once (can reuse tip for type 1 water but not for anything else)

- Tea: Set 100-1000 uL pipette to 0250 and dispense once

Tests- record all in notebook online and paper

- Lamp test (make copy of file) (morning)

- Water Raman SNR (morning)

- Raman Normalize (morning)

- Quinine Sulfate Units (morning)

- Raman Units (morning and evening)

- Evening run only does the “Raman Units” test

Run samples- morning and evening

- Make a copy of the metatable manual, fill out the table

- Make a new method for today’s samples

- Load samples into trays and double check placement

- Make sure the tray and vials are level

- Check if tubing compartment is in aqualog and there is enough type 1 water in the rinse

- Wait to collect blank

Clean up- morning and evening

- Return equipment and samples/standards to their place

- Put the used samples back in the cold room labeled with “ran on aqualog (today’s date)”

- Discard tea standard and rinse with type 1 water 3x. Place upside down on cuvette holder to dry

- Rinse type 1 water container out with type 1 water. Leave standing up on shelf with cap on loosely