638-94599J

Apr. 2021

Total Organic Carbon Analyzer

TOC-LCSH/CSN

User's Manual

Read this manual thoroughly before you use the product.

Keep this manual for future reference.

This page is intentionally left blank.

0

0

0Introduction

Read this Instruction Manual thoroughly before using the product.

Thank you for purchasing this product. This USER'S MANUAL describes how to use this  product in addition to any accessories and options related to this product. Read this manual  thoroughly before using the product and operate the product in accordance with the  instructions in this manual.

Also, keep this manual for future reference.

IMPORTANT

• If the user or usage location changes, ensure that this Instruction Manual is always kept together with the product.

• If this manual or a product warning label is lost or damaged, immediately contact your Shimadzu representative to request a replacement.

• To ensure safe operation, read all Safety Instructions before using the product. • To ensure safe operation, contact your Shimadzu representative if product installation, adjustment, or re-installation (after the product is moved) is required.

Notice

• Information in this manual is subject to change without notice and does not represent a commitment on the  part of the vendor.

• Any errors or omissions which may have occurred in this manual despite the utmost care taken in its  production will be corrected as soon as possible, although not necessarily immediately after detection. • All rights are reserved, including those to reproduce this manual or parts thereof in any form without  permission in writing from Shimadzu Corporation.

• PARAFILM are registered trademarks of American National Can.

• Other company names and product names mentioned in this manual are trademarks or registered trademarks  of their respective companies.

© 2011-2021 Shimadzu Corporation. All rights reserved.

Original version is approved in English.

Indications Used in This Manual

Cautions and Notes are indicated using the following conventions:

Precautionary Conventions Used in this Manual

The following precautionary conventions are used in this manual.

Convention Description

CAUTION Indicates a potentially hazardous situation which, if not avoided, may result in  minor to moderate injury or equipment damage.

Note: Emphasizes additional information that is provided to ensure the proper use of  this product.

Other Conventions Used in this Manual

Instrument Model Names

Descriptions in this documentation pertain to two models of the TOC-L Total Organic  Carbon Analyzer, namely the TOC-LCSH and TOC-LCSN. These models will be referred to  in relevant descriptions as the H Type and N Type, respectively.

Supplementary Conventions

In addition to the precautionary conventions, the "Tip" and "Reference" conventions are  used to supply additional information.

Convention Description

TIP » Indicates a particular technique, alternative means of operation, or useful advice.

Reference: Indicates a reference location in the manual where more detailed information can  be found.

ii TOC-LCSH/CSN

Safety Instructions

To ensure safe product operation, read these important safety instructions carefully before  use and follow all CAUTION instructions given in this section.

Product Applications

This instrument is a total organic carbon analyzer.

Use this instrument ONLY for the intended purpose.

Using this instrument for any other purpose could cause accidents.

Installation Site

Do not install this product at a site where fire is prohibited.

Installation

To ensure safe operation, contact your Shimadzu representative if product installation,  adjustment, or re-installation (after the product is moved) is required.

When lifting the instrument, lift from the bottom and at both sides of the instrument. Do not lift it by the front access door, as this may cause instrument damage.

Operation

The contents of the hard disk of the personal computer may be lost as the result of an  unforeseen accident.

Always create a backup to protect critical data from such accidents.

Do not measure or inject explosive, ignitable, or flammable materials into the instrument.

TOC-L iii CSH/CSN

In an Emergency

In an emergency, turn OFF the main power switch on the left side towards the top of the  TOC-L. Also, turn OFF the power switch on the left side of the ASI-L.

Also, when the product is used again, inspect the product and, if necessary, contact your  Shimadzu representative to request servicing.

Emergency Stop Procedure

1. Close the main valves on the gas cylinders.

2. Turn OFF the main power switch on the left side towards the top of the TOC-L. Also,  turn OFF the power switch on the left side of the ASI-L.

3. Unplug the power cable.

During a Power Outage

Take the following measures in the event of a power outage.

1. Turn OFF the main power switch on the left side towards the top of the TOC-L. If the  main power switch is not turned OFF, the instrument automatically restarts when

power is restored. The power switch on the ASI-L does not need to be turned OFF.

2. After power is restored, turn ON the main power switch on the left side towards the  top of the TOC-L. Confirm that the instrument operates properly before use.

Warning Labels and Residual Risks

In order to ensure safety, warning labels are attached in places requiring caution.

If a warning label is lost or damaged, obtain a new label through your Shimadzu

representative and attach it in the correct position. A residual risk indicates a risk that could  not be reduced or eliminated in the process of design and manufacture. Be sure to

thoroughly read and understand the contents in "Description" before using the product.

iv TOC-LCSH/CSN

1

2

3

TOC-LCSH/CSN v

| No. | Description | Protective Measure |
| --- | --- | --- |
| 1 | High temperature caution Risk of burn | The area near the center of the electric furnace becomes extremely hot when the furnace is heating up and when the furnace has reached operating temperature (680 °C or 720 °C). Do not touch the electric furnace as this could result in burns. Allow the furnace to cool before performing any maintenance on the combustion tube or catalyst. |
| 2 | High temperature caution Risk of burn | Cool down the furnace sufficiently before performing maintenance. |
| 3 | Electric shock caution Risk of electric shock | Do NOT remove the panel. |

vi TOC-LCSH/CSN

4

5

6

| No. | Description | Protective Measure |
| --- | --- | --- |
| 4 | Corrosion hazard Soda lime is a corrosive chemical substance. | Soda lime is a corrosive chemical substance. Be careful not to touch or spill it. For details, see Section 5.10 "Safety Data Sheets" on page 381. |
| 5 | High pressure caution Risk of injury | During operation, the pressure in the CO2 absorber of the gas purifier is about 200 kPa. To prevent injuries, turn off the carrier gas and switch off the power to the TOC main unit. Wait until the pressure inside the container and the instrument flow lines return to normal atmospheric pressure before performing any maintenance. |
| 6 | Corrosion hazard The liquid waste may contain corrosive substances. | Acid and other corrosive substances are present in the liquid waste drained from this instrument. Be careful to not touch or spill the liquid waste. Liquid waste may have a pH level less than 2.0 depending on measurement conditions. Liquid waste with a pH less than 2.0 is classed as industrial waste subject to special control. Request transfer by an authorized industrial waste transporter to an authorized industrial waste processing facility to process any industrial waste subject to special control. If you wish to perform deacidification on your own premises, contact the environmental agency in your jurisdiction. |

TOC-L vii CSH/CSN

7

| No. | Description | Protective Measure |
| --- | --- | --- |
| 7 | Corrosion hazard The liquid waste may contain corrosive substances. | Acid and other corrosive substances are present in the liquid waste drained from this instrument. Be careful to not touch or spill the liquid waste. Liquid waste may have a pH level less than 2.0 depending on measurement conditions. Liquid waste with a pH less than 2.0 is classed as industrial waste subject to special control. Request transfer by an authorized industrial waste transporter to an authorized industrial waste processing facility to process any industrial waste subject to special control. If you wish to perform deacidification on your own premises, contact the environmental agency in your jurisdiction. |

viii TOC-LCSH/CSN

8

9

0

TOC-LCSH/CSN ix

| No. | Description | Protective Measure |
| --- | --- | --- |
| 8 | Corrosion hazard The liquid waste may contain corrosive substances. | There is a risk of corrosive substance leakage. Securely connect the Y-shaped fitting to the tip of the tubing. Be careful not to contact or spill the liquid. |
| 9 | Corrosion hazard The IC reaction solution is a corrosive chemical substance. | The IC reaction solution contains corrosive (acidic) phosphoric acid. Be careful not to touch or spill it. For details, see Section 5.10 "Safety Data Sheets" on page 381. |
| 0 | Corrosion hazard Hydrochloric acid is a corrosive chemical substance. | Hydrochloric acid is a corrosive chemical substance. Be careful not to touch or spill it. For details, see Section 5.10 "Safety Data Sheets" on page 381. |

x TOC-LCSH/CSN

Warranty

Shimadzu provides the following warranty for this product.

Period:

Please contact your Shimadzu representative for information about the period of this  warranty.

Description:

If a product/part failure occurs for reasons attributable to Shimadzu during the warranty  period, Shimadzu will repair or replace the product/part free of charge. However, in the  case of products which are usually available on the market only for a short time, such as

personal computers and their peripherals/parts, Shimadzu may not be able to provide  identical replacement products.

Limitation of Liability:

(1) In no event will Shimadzu be liable for any lost revenue, profit or data, or for special,  indirect, consequential, incidental or punitive damages, however caused regardless of  the theory of liability, arising out of or related to the use of or inability to use the

product, even if Shimadzu has been advised of the possibility of such damage.

(2) In no event will Shimadzu's liability to you, whether in contract, tort (including  negligence), or otherwise, exceed the amount you paid for the product.

Exceptions:

Failures caused by the following are excluded from the warranty, even if they occur during  the warranty period.

1) Improper product handling

2) Repairs or modifications performed by parties other than Shimadzu or Shimadzu

designated companies

3) Product use in combination with hardware or software other than that designated by  Shimadzu

4) Computer viruses leading to device failures and damage to data and software,

including the product's basic software

5) Power failures, including power outages and sudden voltage drops, leading to device  failures and damage to data and software, including the product's basic software

6) Turning OFF the product without following the proper shutdown procedure leading  to device failures and damage to data and software, including the product's basic

software

7) Reasons unrelated to the product itself

8) Product use in harsh environments, such as those subject to high temperatures or

humidity levels, corrosive gases, or strong vibrations

TOC-L xi CSH/CSN

9) Fires, earthquakes, or any other act of nature, contamination by radioactive or

hazardous substances, or any other force majeure event, including wars, riots, and

crimes

10) Product movement or transportation after installation

11) Consumable items

Note: Recording media such as floppy disks and CD-ROMs are considered

consumable items.

* If there is a document such as a warranty provided with the product, or there is a  separate contract agreed upon that includes warranty conditions, the provisions of those  documents shall apply.

xii TOC-LCSH/CSN

After-Sales Service and Availability of Replacement Parts

After-Sales Service

If any problem occurs with this product, perform an inspection and take appropriate  corrective action as described in this manual's troubleshooting section. If the problem  persists, or the symptoms are not covered in the troubleshooting section, contact your  Shimadzu representative.

Replacement Parts Availability

Replacement parts for this product will be available for a period of seven (7) years after the  product is discontinued. Thereafter, such parts may cease to be available.

If Shimadzu receives notice of the discontinuation of units or parts, the necessary quantity  for the above period is immediately calculated and secured. However, such units or parts  may cease to be available within seven years after the discontinuation of the product,  depending on the conditions of individual manufacturers and on changes in the quantity  required.

Maintenance, Inspections, and Adjustment

In order to maintain the instrument's performance and obtain accurate measurement data,  daily inspection and periodic inspection are necessary.

• For daily maintenance and periodic inspection, see "4 Maintenance".

• For consumable parts and maintenance parts, see "5 Reference Materials".

The replacement cycle for each part is given at a rough estimate.

Replacement may be required earlier than the replacement cycle depending on the

environment and frequency of usage.

TOC-L xiii CSH/CSN

Disposal Precautions

• Dispose of this instrument according to local regulations for the relevant area and  request processing by an industrial waste disposal company. Gather and seal insulation  material waste into plastic bags and request disposal by an industrial waste disposal  company licensed to process glass, concrete, and ceramic waste.

• All used catalyst should be sent to a licensed industrial waste disposal company.

• All used CO2 absorbers should be sent to a licensed industrial waste disposal company. • All used halogen scrubbers should be sent to a licensed industrial waste disposal  company.

• All used catalyst from the ozone treatment unit should be sent to a licensed industrial  waste disposal company.

• The optional measurement circuit for purgeable organic carbon (POC) uses a CO2 absorber filled with lithium hydroxide to remove carbon dioxide. All used lithium

hydroxide crystals should be sent to a licensed industrial waste disposal company.

• Wear protective glasses and gloves when handling quartz wool.

Inhaling it may cause health impairment.

CAUTION

For California, USA Only

This product contains a battery that contains perchlorate material.

Perchlorate Material - special handling may apply.

See www.dtsc.ca.gov/hazardouswaste/perchlorate

xiv TOC-LCSH/CSN

Electromagnetic Compatibility

This instrument complies with European standard EN50011 class A for electromagnetic  emission and EN61326-1 for electromagnetic immunity.

EN55011 Emissions (Electromagnetic Interference)

This instrument complies with European EMI (Electro Magnetic Interference) standard  EN55011 class A equipment. This instrument can be used in any environment except in the  home.

Note: When an electromagnetic disturbance occurs due to nearby this instrument, provide an appropriate distance between instruments to eliminate the disturbance.

In addition, please connect this instrument to a large power supply of the capacity as much as possible.

EN61326-1 Immunity (Electromagnetic Susceptibility)

This instrument complies with European standard EN61326-1 for electromagnetic

immunity.

Stated below are the test conditions.

| IEC 61000-4-2 Electrostatic Discharge | Air: 8 kV, Direct: 4 kV |
| --- | --- |
| IEC 61000-4-3 Radiated, Radio-Frequency, Electromagnetic Field | 10 V/m (0.08-1 GHz), 3 V/m (1.4-2 GHz), 1 V/m (2-2.7 GHz) |
| IEC 61000-4-4 Transient/Burst (Electrical Fast Transients) | 2.0 kV to AC power line and ground |
| IEC 61000-4-5 Voltage Surge | 1.0 kV line to line, 2 kV line to ground |
| IEC 61000-4-6 Conducted RF Immunity | 3 Vrms (150 kHz-80 MHz) |
| IEC 61000-4-11 Voltage Variations/Dips/Interrupts | 0 % UT 1 cycles (20 ms), 70 % UT 25 cycles (500 ms), 0 % UT 250 cycles (5000 ms) |

Note: Compliance to the standard does not ensure that the instrument can work with any level of Electromagnetic interference stronger than the level tested.

Interference greater than the value specified in the condition above may cause malfunction of the instrument.

To above electromagnetic disturbances, follow the recommendations below.

(1) Before touching the instrument, discharge the electro static charged in operator's body  to ground touching metallic structure connected to ground.

(2) Do not install this instrument in such environment where strong electromagnetic  fields are generated near by.

TOC-LCSH/CSN xv

Contents of This Manual

Chapter 1 - Overview

This chapter provides an overview of the instrument, as well as safety precautions to be  taken while using the instrument.

Chapter 2 - System Description

This chapter identifies and describes the system components and provides a brief

description of optional accessories.

Chapter 3 - Operation

This chapter provides detailed descriptions of the instrument functions, typical preparation  procedures, basic operation procedures such as calibration curve generation, and

application operations such as using optional features.

Chapter 4 - Maintenance

This chapter details the typical maintenance items and periodic checks required to ensure  that the instrument provides consistently high accuracy analyses. Also included is a  description of the error messages and troubleshooting procedures.

Chapter 5 - Reference Information

This chapter provides a detailed description of the analysis principles on which the

instrument is based, specifications, standard and optional accessories, the various screens  displayed on the LCD, as well as the specialized terms associated with this instrument. The  installation procedure is also provided in the event that instrument is moved to another  installation site.

xvi TOC-LCSH/CSN

Table of Contents

Introduction................................................................................................................ i Indications Used in This Manual.................................................................................... ii Precautionary Conventions Used in this Manual................................................... ii Other Conventions Used in this Manual ................................................................ ii

Safety Instructions.......................................................................................................... iii Product Applications.............................................................................................. iii Installation Site....................................................................................................... iii Installation.............................................................................................................. iii Operation................................................................................................................ iii In an Emergency .................................................................................................... iv During a Power Outage.......................................................................................... iv Warning Labels and Residual Risks ...................................................................... iv

Warranty......................................................................................................................... xi After-Sales Service and Availability of Replacement Parts........................................... xiii After-Sales Service................................................................................................. xiii Replacement Parts Availability.............................................................................. xiii Maintenance, Inspections, and Adjustment.................................................................... xiii Disposal Precautions ...................................................................................................... xiv Electromagnetic Compatibility....................................................................................... xv Contents of This Manual ................................................................................................ xvi

1 1Overview ................................................................................................................... 1 1.1 Overview ........................................................................................................................ 2 1.2 Operational Precautions ................................................................................................. 3

2 2System Description................................................................................................... 7 2.1 TOC-LCSH/CSN Construction........................................................................................... 8 2.1.1 Front View ............................................................................................................. 8 2.1.2 Right Side View ..................................................................................................... 8 2.1.3 Left Side View ....................................................................................................... 9

2.1.4 Rear View............................................................................................................... 10 2.1.5 Front Internal View (Left Diagonal)...................................................................... 11 2.1.6 Top Internal View (Left Diagonal) ........................................................................ 12 2.1.7 Keyboard Layout.................................................................................................... 12 2.1.8 Flow Diagrams....................................................................................................... 14

2.2 Autosampler (Option)..................................................................................................... 19 2.2.1 Front View (Left Diagonal) ................................................................................... 19 2.2.2 Rear View (Left Diagonal)..................................................................................... 19

2.3 8-Port Sampler OCT-L (Option) .................................................................................... 20 2.3.1 Front View (Left Diagonal) ................................................................................... 20 2.3.2 Rear View (Right Diagonal) .................................................................................. 20

2.4 TN Unit (Option)............................................................................................................ 21 2.4.1 Front View (Right Diagonal) ................................................................................. 21 2.4.2 View Inside the Door on the Right Side ................................................................ 21 2.4.3 Rear View............................................................................................................... 22

TOC-L xvii CSH/CSN

Table of Contents

3 3Operation.................................................................................................................. 23 3.1 Analysis Preparation....................................................................................................... 24 3.1.1 Catalyst Filling ....................................................................................................... 24 3.1.1.1 Catalyst for TOC Analysis........................................................................... 25 3.1.1.2 TOC/TN Catalyst (Option) .......................................................................... 27 3.1.2 Installing the Cooling Tube Unit (H Type Only) ................................................... 29 3.1.3 Connecting the Combustion Tube.......................................................................... 30 3.1.4 Connecting the Halogen Scrubber.......................................................................... 39 3.1.5 Installing the Syringe.............................................................................................. 40 3.1.6 Water Supply to the Dehumidifier Drain Vessel.................................................... 42 3.1.7 Filling the Humidifier with Water.......................................................................... 43 3.1.8 Preparation of Dilution Water ................................................................................ 44 3.1.9 Acid Preparation..................................................................................................... 45 3.1.10 Installation of the CO2 Absorber ............................................................................ 46 3.1.11 Printer Preparation.................................................................................................. 46 3.1.12 Installing the Suspended Solids Kit........................................................................ 47 3.1.13 IC Analysis Preparations ........................................................................................ 49 3.1.14 Preparation for POC Analysis ................................................................................ 51 3.1.15 TN Analysis Preparations (When Using the Optional TN Unit)............................ 54 3.1.15.1 NOx Absorber .............................................................................................. 54 3.1.16 B-Type Halogen Scrubber

(When Using the Optional B-Type Halogen Scrubber) ......................................... 55 3.1.17 Instrument Parameters............................................................................................ 57 3.1.17.1 Buzzer .......................................................................................................... 57 3.1.17.2 Date and Time.............................................................................................. 58 3.1.17.3 Catalyst Type ............................................................................................... 60 3.1.17.4 Concentration Unit....................................................................................... 61 3.1.17.5 Printer Setup................................................................................................. 62 3.1.17.6 Suspended Solids Analysis .......................................................................... 64 3.1.17.7 Syringe Washes............................................................................................ 64 3.1.17.8 Automatically Change the Injection Volume............................................... 66 3.1.17.9 Automatically Change Dilution Factor ........................................................ 67 3.1.17.10 Multiple Injections ....................................................................................... 68 3.1.17.11 Setting Conversion Equation........................................................................ 69 3.1.17.12 Automatic Regeneration of IC Reagent (H Type) ....................................... 70 3.1.17.13 Analysis Span............................................................................................... 72 3.1.17.14 Counter Limit Values................................................................................... 73 3.1.17.15 Screen Saver................................................................................................. 74 3.1.18 Preparation and Storage of Standard Solutions...................................................... 75 3.1.18.1 Preparation of Standard Solutions................................................................ 75 3.1.18.2 Storage of Standard Solutions...................................................................... 77 3.1.19 Sample Preparation................................................................................................. 78 3.1.19.1 High Sensitivity Analysis............................................................................. 78 3.1.19.2 TC Blank Check Analysis............................................................................ 81 3.1.19.3 Pretreatment for IC Removal (Sparging) ..................................................... 83 3.1.19.4 Analysis of Samples Containing Suspended Solids..................................... 84 3.1.19.5 Handling Samples Containing Acids, Bases or Salts................................... 85 3.2 Analysis .......................................................................................................................... 87 3.2.1 Starting Up the Instrument ..................................................................................... 87

xviii TOC-LCSH/CSN

Table of Contents

3.2.1.1 Turning the Power ON................................................................................. 87 3.2.1.2 Setting the Carrier Gas Supply Pressure...................................................... 88 3.2.1.3 Setting the Carrier Gas Supply Pressure

(When Using the Carrier Gas Regulator Option) ........................................ 88

3.2.1.4 Setting the Flow Rate of Carrier Gas........................................................... 89 3.2.1.5 Setting the Sparge Gas Flow Rate

(When Using the Optional External Sparge Kit)......................................... 91

3.2.1.6 Electric Furnace ........................................................................................... 92 3.2.1.7 Turning On Power to the TNM-L (Option)................................................. 93 3.2.1.8 Setting the TNM-L Ozone Air Flow Rate (Option) .................................... 93

3.2.2 Calibration Curve Analysis.................................................................................... 94 3.2.2.1 Single-point Calibration Curve.................................................................... 94 3.2.2.2 Multi-point Calibration Curve Analysis ...................................................... 101 3.2.2.3 Multi-point Calibration Curve Analysis with Diluted Solution................... 105 3.2.2.4 Changing Parameters for a Generated Calibration Curve ........................... 113

3.2.3 Sample Analysis.................................................................................................... 114 3.2.3.1 Single-mode Analysis.................................................................................. 114 3.2.3.2 Multi-mode Analysis ................................................................................... 123 3.2.3.3 Multi-mode Analysis - Simultaneous Analysis ........................................... 126 3.2.3.4 Multi-mode Analysis - Common Injection Analysis................................... 129

3.2.4 Creating a Method.................................................................................................. 134 3.2.5 Stopping Analysis .................................................................................................. 137 3.2.6 Discarding Data...................................................................................................... 138 3.2.7 Ending Analysis ..................................................................................................... 139

3.3 Analysis with the Autosampler ...................................................................................... 142 3.3.1 Sample Preparation for Autosampler Measurement .............................................. 142 3.3.1.1 Vial Types.................................................................................................... 142 3.3.1.2 Filling Vials with Sample ............................................................................ 144 3.3.1.3 Placing Vials in the Vial Rack..................................................................... 145 3.3.1.4 Rinse Water Preparation .............................................................................. 146 3.3.1.5 High Sensitivity ASI-L Analysis ................................................................. 147 3.3.1.6 Acid Addition in NPOC Analysis................................................................ 148 3.3.1.7 Sparging in NPOC Analysis ........................................................................ 148 3.3.1.8 Stirrer Preparation........................................................................................ 149 3.3.2 Setting General Measurement Conditions ............................................................. 151 3.3.3 Setting Analysis Parameters................................................................................... 154 3.3.3.1 Calibration Curve Analysis.......................................................................... 154 3.3.3.2 Single-point Calibration Curve Analysis..................................................... 155 3.3.3.3 Multi-point Calibration Curve Analysis ...................................................... 159 3.3.3.4 Multi-point Calibration Curve Analysis with Diluted Solution................... 161 3.3.3.5 Sample Analysis Settings............................................................................. 163 3.3.3.6 Single-mode Analysis.................................................................................. 165 3.3.3.7 Multi-mode Analysis ................................................................................... 168 3.3.3.8 Multi-mode Analysis - Simultaneous Analysis ........................................... 170 3.3.3.9 Multi-mode Analysis - Common Injection Analysis................................... 172 3.3.3.10 Editing the ASI Schedule Settings............................................................... 175 3.3.4 Analysis.................................................................................................................. 178 3.3.4.1 Starting Analysis.......................................................................................... 178 3.3.4.2 Interrupting Analysis ................................................................................... 180 3.4 Analysis with the 8-Port Sampler................................................................................... 181

TOC-LCSH/CSN xix

Table of Contents

3.4.1 Preparation for Sample Analysis Using the OCT-L 8-Port Sampler ..................... 181 3.4.1.1 Sample Container......................................................................................... 181 3.4.1.2 Loading the Sample...................................................................................... 181 3.4.1.3 High Sensitivity Using the OCT-L .............................................................. 182

3.4.2 Setting Analysis Parameters................................................................................... 183 3.4.2.1 Calibration Curve Analysis.......................................................................... 183 3.4.2.2 Single-point Calibration Curve Analysis ..................................................... 184 3.4.2.3 Multi-point Calibration Curve Analysis....................................................... 188 3.4.2.4 Multi-point Calibration Curve Analysis with Diluted Solution................... 190 3.4.2.5 Sample Analysis Settings............................................................................. 193 3.4.2.6 Single-mode Analysis .................................................................................. 193 3.4.2.7 Multi-mode Analysis.................................................................................... 197 3.4.2.8 Multi-mode Analysis - Simultaneous Analysis............................................ 198 3.4.2.9 Multi-mode Analysis - Common Injection Analysis ................................... 201 3.4.2.10 Editing the ASI Schedule Settings ............................................................... 203

3.4.3 Analysis .................................................................................................................. 206 3.4.3.1 Starting Analysis .......................................................................................... 206 3.4.3.2 Interrupting Analysis.................................................................................... 208

3.5 Data Review.................................................................................................................... 209 3.5.1 Reviewing Calibration Curve List.......................................................................... 209 3.5.2 Reviewing the Method List .................................................................................... 213 3.5.3 Reviewing Data Reports......................................................................................... 215

3.6 Outputting Data to a USB Memory Device.................................................................... 218 3.7 Applications Operations (H Type Only)......................................................................... 221 3.7.1 Manual Injection Analysis...................................................................................... 221 3.7.2 Gas Sample Analysis (When Using the Optional Manual Injection Kit)............... 229

4 4 Maintenance ............................................................................................................ 235 4.1 Daily Inspection.............................................................................................................. 236 4.1.1 Checking Dilution Water........................................................................................ 236 4.1.2 Checking Acid........................................................................................................ 236 4.1.3 Checking Drain Vessel Water Level ...................................................................... 236 4.1.4 Checking Humidifier Water Level ......................................................................... 237 4.1.5 Inspecting the Liquid in the B-Type Halogen Scrubber (Option).......................... 237 4.2 Periodic Inspections........................................................................................................ 238 4.2.1 Catalyst Regeneration............................................................................................. 238 4.2.2 Washing or Replacing Catalyst .............................................................................. 239 4.2.3 Replacing the Carrier Gas Purification Tube and Catalyst .................................... 242 4.2.4 Washing / Replacing the Combustion and Carrier Gas Purification Tubes ........... 244 4.2.5 Replacing the High-Purity Air Cylinder................................................................. 244 4.2.6 Replacing the CO2 Absorber .................................................................................. 244 4.2.7 Replacing the Halogen Scrubber............................................................................ 245 4.2.8 Syringe Replacement.............................................................................................. 247 4.2.9 Replacing the Syringe Plunger Tip......................................................................... 248 4.2.10 Replacing The Sliding Sample Injector O-Rings ................................................... 250 4.2.11 Replacing the Syringe Pump 8-port Valve Rotor................................................... 250 4.2.12 Replacing the CO2 Absorber for the Optional Carrier Gas Purifier....................... 253 4.2.13 Replenishing IC Reaction Solution ........................................................................ 254

xx TOC-LCSH/CSN

Table of Contents

4.2.14 Replacing the Pump Head of the IC Reaction Solution Pump............................... 255 4.2.15 Replacing the Filler of the Optional POC Analysis CO2 Absorber ....................... 256 4.2.16 Replacing the Liquid in the B-Type Halogen Scrubber (Option).......................... 256

4.3 Autosampler (ASI-L) Maintenance................................................................................ 257 4.3.1 Rinse Bottle Inspection .......................................................................................... 257 4.3.2 Sample Catcher Inspection..................................................................................... 257 4.3.3 Replacing the Pump Head of the Rinse Pump ....................................................... 258

4.4 8-Port Sampler (OCT-L) Maintenance........................................................................... 259 4.4.1 Replacing the 8-Port Valve Rotor.......................................................................... 259 4.5 TN Unit (TNM-L) Maintenance..................................................................................... 261 4.5.1 Replacing the Catalyst in the Ozone Treatment Unit............................................. 261 4.5.2 Inspecting the Chemiluminescence Detector

and Replacing the Ozone Generator ...................................................................... 263 4.5.3 Replacing the NOx Absorber................................................................................. 263 4.6 Troubleshooting.............................................................................................................. 264 4.6.1 Error Messages....................................................................................................... 264 4.6.2 Troubleshooting ..................................................................................................... 270 4.6.2.1 TOC-L.......................................................................................................... 270 4.6.2.2 ASI-L ........................................................................................................... 275 4.6.2.3 8-Port Sampler OCT-L ................................................................................ 276 4.6.2.4 Remedial Actions for Poor Repeatability .................................................... 277 4.6.2.5 Corrective Action for Combustion Tube Breakage

Inside the Electric Furnace .......................................................................... 279

5 5 Reference Materials ................................................................................................ 281 5.1 Principles of Analysis..................................................................................................... 282 5.1.1 Principle of TC (Total Carbon) Analysis............................................................... 282 5.1.2 Principles of IC (Inorganic Carbon) Analysis........................................................ 283 5.1.3 Principle of NPOC (Non-Purgeable Organic Carbon) Analysis............................ 284 5.1.4 Principle of POC (Purgeable Organic Carbon) Analysis....................................... 285 5.1.5 Principles of Measuring TOC ................................................................................ 286 5.1.6 Principles of Measuring TN (Total Nitrogen)........................................................ 286 5.2 Analysis-Related Technical Information........................................................................ 287 5.2.1 Peak Area Analysis ................................................................................................ 287 5.2.2 Peak Shape ............................................................................................................. 287 5.2.3 Calibration Curves ................................................................................................. 288 5.2.3.1 Types of Calibration Curves........................................................................ 288 5.2.3.2 Shifting of Calibration Curves..................................................................... 289 5.2.4 Sparging During Standard Solution Analysis ........................................................ 290 5.2.5 Automatic Selection of the Optimal Calibration Curve......................................... 290 5.2.6 Automatically Changing the Injection Volume ..................................................... 292 5.2.7 Automatically Changing the Dilution Factor......................................................... 292 5.2.8 Multiple Injections ................................................................................................. 293 5.3 Screen Descriptions........................................................................................................ 294 5.3.1 "Initial Display" ..................................................................................................... 294 5.3.2 "Background Monitor" Screen............................................................................... 295 5.3.3 "Conditions" Screen ............................................................................................... 296

TOC-L xxi CSH/CSN

Table of Contents

5.3.4 "Maintenance" Screen ............................................................................................ 301 5.3.5 "Data Report - Measurement Log" Screen............................................................. 303 5.3.6 "Data Report - Calibration Curves" Screen............................................................ 305 5.3.7 "Calibration Settings" Screen ................................................................................. 308 5.3.8 Calibration Curve Generation................................................................................. 311

5.3.8.1 "Calibration" Screen..................................................................................... 311 5.3.8.2 During Analysis ........................................................................................... 312 5.3.8.3 Completion of Analysis Point ...................................................................... 314 5.3.8.4 Analysis Completion.................................................................................... 314 5.3.8.5 Calibration Curve Protection........................................................................ 316

5.3.9 "Measurement Parameters" Screen ........................................................................ 317 5.3.9.1 Calibration Curve List.................................................................................. 321 5.3.10 Analysis .................................................................................................................. 322 5.3.10.1 During Analysis ........................................................................................... 324 5.3.10.2 Viewing the Analysis Parameters ................................................................ 325 5.3.10.3 Viewing the Ready Parameters .................................................................... 326 5.3.10.4 Analysis Completion.................................................................................... 327 5.3.11 "Method Store" Screen ........................................................................................... 328 5.3.11.1 Loading a Method ........................................................................................ 329 5.3.12 "ASI Schedule - Settings" Screen........................................................................... 330 5.3.12.1 "ASI Calibration Settings" Screen ............................................................... 332 5.3.12.2 "ASI Measurement Settings" Screen ........................................................... 333 5.3.12.3 ASI Schedule - Starting Analysis................................................................. 334 5.3.12.4 ASI Schedule - During Analysis.................................................................. 336 5.3.12.5 ASI Schedule - Viewing the Schedule......................................................... 337 5.3.12.6 ASI Schedule - Viewing the Calibration Curve Generation Parameters ..... 338 5.3.12.7 ASI Schedule - Viewing Analysis Parameters............................................. 339 5.3.12.8 ASI Schedule - Viewing the Ready Parameters........................................... 340 5.4 Specifications.................................................................................................................. 341 5.4.1 TOC-LCSH/CSN ........................................................................................................ 341 5.4.2 ASI-L...................................................................................................................... 342 5.4.3 OCT-L 8-Port Sampler........................................................................................... 342 5.4.4 TNM-1.................................................................................................................... 343 5.4.5 POC Measurement Kit ........................................................................................... 343 5.4.6 Carrier Gas Purifier Kit .......................................................................................... 343 5.4.7 Manual Injection Kit (Gas Sample Injection Kit) .................................................. 344 5.4.8 B-Type Halogen Scrubber...................................................................................... 344 5.4.9 External Sparge Kit ................................................................................................ 344 5.4.10 LAN Output............................................................................................................ 345 5.4.10.1 Communication Specifications .................................................................... 345 5.4.10.2 Connection Cable......................................................................................... 345 5.4.10.3 Transmission Content................................................................................... 345 5.5 Standard Accessories...................................................................................................... 351 5.6 Special Accessories ........................................................................................................ 353 5.7 Consumable Parts List .................................................................................................... 355 5.7.1 Consumable Parts List............................................................................................ 355 5.7.2 Recommended Consumable Part Replacement Period .......................................... 358 5.8 Maintenance Parts List ................................................................................................... 359

xxii TOC-LCSH/CSN

Table of Contents

5.9 Installation...................................................................................................................... 361 5.9.1 Before Installation.................................................................................................. 361 5.9.1.1 Parts Inspection............................................................................................ 361 5.9.2 Installation Site....................................................................................................... 361 5.9.2.1 Installation Site Selection ............................................................................ 361 5.9.2.2 Installation Site Conditions.......................................................................... 363 5.9.3 Installation Procedure............................................................................................. 364 5.9.3.1 Connecting Power Supply and Ground ....................................................... 364 5.9.3.2 Changing the Power Supply Voltage........................................................... 365 5.9.3.3 Connecting Gas............................................................................................ 365 5.9.3.4 Installing the Drain Container...................................................................... 368 5.9.4 Installing the ASI-L................................................................................................ 369 5.9.5 OCT-L Installation................................................................................................. 375 5.9.6 Installing the TNM-L ............................................................................................. 378 5.9.6.1 TNM-L TN Unit Installation ....................................................................... 378 5.9.6.2 Connecting Ozone Source Air ..................................................................... 378 5.9.6.3 Connecting Gas Exhaust Tubing ................................................................. 380 5.10 Safety Data Sheets.......................................................................................................... 381 5.10.1 Hydrochloric Acid (HCl) ....................................................................................... 381 5.10.2 Phosphoric Acid..................................................................................................... 384 5.10.3 IC Reagent.............................................................................................................. 387 5.10.4 CO2 Absorber, Ca(OH)2, KOH, NaOH: Soda Lime.............................................. 390 5.10.5 Ceramic Fiber......................................................................................................... 393 5.10.6 Catalyst Platinum Black......................................................................................... 395 5.10.7 NH3 Absorbent ...................................................................................................... 399 5.10.8 Silica Glass Wool................................................................................................... 402

TOC-L xxiii CSH/CSN

Table of Contents

This page is intentionally left blank.

xxiv TOC-LCSH/CSN

1

1Overview

This chapter provides an overview of the instrument and the operational precautions for  instrument use.

1.1 Overview

Provides and overview of this instrument.

1.2 Operational Precautions

Details precautions necessary to ensure that the instrument is used correctly.

1

1

1.1 Overview

1.1 Overview

The TOC-LCSH/TOC-LCSN (hereinafter referred to as TOC-L) is an instrument that  measures the amount of total carbon (TC), inorganic carbon (IC) and total organic carbon  (TOC) in water. Purgeable organic carbon (POC) can also be measured when that accessory  is installed. "Oxidative combustion-infrared analysis" is a widely used TOC measurement  method that has been adopted by the JIS1 and other international standards. By installing  the optional TN unit, it is also possible to measure total water-borne nitrogen (TN) using  the principles of "oxidative combustion-chemiluminescence."

Moreover, when the ASI-L Autosampler or OCT-L 8-Port Sampler is used in combination  with the main unit, a fully automatic system is created, allowing automatic analysis of  multiple samples.

This instrument is a qualified product in Shimadzu's ecolabel program.

The power consumption of this instrument has been reduced by 36 % for a

supply voltage of 100 V and 43 % for a supply voltage of 200 V compared

to previous TOC-V models. (Usage conditions: (8 hours of operation per

day) ⋅ (5 hours of operation per week))

CAUTION

Safety may be compromised if the instrument is used in a fashion other than that

indicated in this USER'S MANUAL. Pay particular attention to the CAUTION

sections to ensure safety when using the instrument.

1. JIS K-0102 "Industrial Waste Water Testing", JIS K-0551: "Total organic carbon (TOC) testing methods  for ultra-pure water", U.S.Pharmacopoeia, EPA 415.1, EPA 9060A, ASTM D7573-09, Standard Methods  for Examination of Water and Waste Water 5310B

2 TOC-LCSH/CSN

1.2 Operational Precautions

1.2 Operational Precautions

Take the following precautions when using this instrument.

CAUTION

• This instrument is used for measuring the amount of organic and nitrogenous substances  in water (additional options allow gaseous and solid sample measurement). Do NOT  measure liquids other than water samples. <Part damage prevention>

• Do not touch the electric furnace while it is heating. The center of the electric furnace  (near the combustion tube insertion opening) reaches very high temperatures, and burns  may result. <Burn prevention>

• To prevent high sensitivity catalyst from being forced upward in the combustion tube,  verify that carrier gas is flowing when heating the electric furnace. <Prevents the

floating of high sensitivity catalyst>

• Allow the electric furnace to cool to room temperature for at least six hours after  switching off the power supply by shutting down the instrument before removing or  exchanging the combustion tube. Burns may result if this procedure is attempted when  the furnace is at a high temperature. <Burn prevention>

• Do not heat the electric furnace without the combustion tube installed. The parts of the  TC injection port will be exposed to radiant heat from the center of the electric furnace  and may become deformed. If the furnace must be heated without the combustion tube  installed, either remove the TC injection port parts from the instrument in advance to

prevent deformation or cover the furnace core opening with quartz wool or another  suitable insulator. <Part damage prevention>

• To reduce back pressure, verify that the external drain tubing connected to the waste  port on the right side of the instrument does not touch the surface of the liquid in the  waste container. The height of the external tube must always be lower than the height of  the drain port. Note that excessive back pressure will prevent sufficient drainage and  cause overflowing from the drain pot. <Part corrosion prevention>

TOC-L Figure 1.1

Drain discharge tubing (Avoid  installing the tubing as shown  by the dotted line.)

Drain container (Prepare a  suitable drain container.)

TOC-LCSH/CSN 3

1.2 Operational Precautions

• Hand tighten the various connectors to the 8-port valve. Use of a tool may cause  deformation and leakage in the valve. <Part damage prevention>

• Stop operation of the instrument when performing maintenance on any of the drive  components, such as connecting lines to the 8-port valve, or connecting/ disconnecting  lines on the sample injector. If these parts are manipulated or if the lines are

disconnected during operation, the drive component could move, causing injury or

discharge of liquid. <Injury prevention, corrosion prevention>

• Specify the Shut down instrument option in the Shutdown window to shut down the  instrument. The electric furnace turns OFF and the power supply shuts off after 30

minutes except for a minimum amount of power necessary to keep the software

running. This shutdown process extends the life of the TC injection port parts by

preventing unnecessary wear. This procedure does not apply to emergency shutdown.  <Maintaining durability of TC injection port parts>

• Turn the lever of the 3-way valve at the bottom of the IC reaction vessel to the left  position. The lower position allows carrier gas to flow at a pressure of 200kPa upstream  from the 3-way valve, causing possible disconnection of joints and damage to parts such  as the combustion tube.

Carrier gas is discharged to the front of the IC reaction vessel. (Turn the lever to this position only when detaching an IC reaction vessel that contains liquid.)

Never turn the lever to this position.

Figure 1.2 Lever Position of 3-Way Valve on the IC Reaction Vessel

• Always verify that the ozone air source is supplying air at a rate of 0.5 L/min when  using the TNM-L. This helps to prevent ozone leaks and line corrosion. <Ozone leak  prevention, line corrosion prevention>

• If an ozone odor is detected while using the TNM-L, immediately stop the current  measurement (without turning off the gas supply) and contact your Shimadzu  representative to request maintenance. Ozone can cause respiratory organ injury and  equipment damage. <Prevention of respiratory organ injury and equipment damage>

• The optional POC measurement circuit incorporates a carbon dioxide (CO2) absorber  packed with lithium hydroxide for removal of carbon dioxide gas. Misuse of lithium  hydroxide can cause bodily harm. During packing or replacing the CO2 absorber  container, exercise adequate precaution. After use, treat the material as industrial waste.  The properties and handling guidelines for lithium hydroxide are described in the  material safety data sheets (MSDS) in Section 5.10 "Safety Data Sheets" on page 381.

4 TOC-LCSH/CSN

1.2 Operational Precautions

• To clean the instrument, wipe the surface with a soft cloth moistened with water. Do not  use chemicals.

• Keep samples and chemicals away from the PC and printer.

• Repairs to the inside of this instrument can be dangerous. Contact your Shimadzu  representative to have a trained service engineer perform any repairs.

• Do not disassemble or modify this instrument beyond the scope of the maintenance  procedures in this manual as safety may be compromised.

• Do not use the instrument in hazardous areas (where fire is prohibited) because the  construction of the instrument is not explosion-proof. Also, the instrument cannot be  used in cleanroom environments.

• Select an installation site with a stable ambient temperature and clean surroundings free  from corrosive gas, organic gas, and dust. Dirty surroundings in particular can cause  measurement errors when performing high sensitivity measurements. Install the

instrument on a sturdy, horizontal bench free from vibrations and physical impacts.

• In order to obtain favorable measurement values, do not install the instrument near  windows subject to direct sunlight, other sources of heat, or locations directly subject to  cool air from air conditioning, or near equipment that generates strong magnetic fields,  electric fields, or high frequency waves.

• Use the instrument with a grounded power supply of AC 100 ± 10 V.

• Do not use the following characters when naming files or systems to be saved:

/ , . \ (back slash) : ; (these are all normal-width characters), space (normal width, em  width)

• Hand-tighten the connectors to the 8-port valve of the OCT-L. Over-tightening with a  tool may exert excessive force to the internal valve body, causing deformation and

internal leakage. <Part damage prevention>

• Shut down the instrument before connecting or disconnecting tubing to the 8-port valve  or performing maintenance to the drive parts. Liquid may be discharged from tubes  during operation. <Part corrosion prevention>

• Do not write on the periphery of the vial rack or stick adhesive tape onto the vial rack. If  aluminum foil or water droplets are adhered to the periphery of the vial rack, remove  them. These may prevent the vial rack from rotating correctly.

• At the start of measurement, the ASI-L rotates the vial rack by one revolution to detect  its home position and performs initialization. In order to perform measurement

correctly, the ASI-L may also detect shift in the vial rack position during measurement  and perform correction by rotating the vial rack by one revolution. This does not mean  that an abnormality has occurred.

TOC-LCSH/CSN 5

1.2 Operational Precautions

This page is intentionally left blank.

6 TOC-LCSH/CSN

2

2System Description

This chapter identifies the components of the TOC-LCSH/CSN, autosampler and the TN  unit.

2.1 TOC-LCSH/CSN Construction

This section shows the construction of the TOC-L main unit using front, right,  left, rear, and internal front and top views depicted diagonally from the left of the  instrument, as well as flow diagrams.

2.2 Autosampler (Option)

This section shows the construction of the ASI-L unit using front and rear views  depicted diagonally from the left.

2.3 8-Port Sampler OCT-L (Option)

This section shows the construction of the OCT-L unit using front and rear views  depicted diagonally from the left and right.

2.4 TN Unit (Option)

This section shows the construction of the TNM-L unit using a front view  depicted diagonally from the right, a right view showing inside the door, and a  rear view.

2

2

2.1 TOC-LCSH/CSN Construction

2.1 TOC-LCSH/CSN Construction

2.1.1 Front View

Power switch (with

LED)

Indicators

Front door

Figure 2.1 Front View

2.1.2 Right Side View

Main power switch

Front door

Figure 2.2 Right Side View

8 TOC-LCSH/CSN

2.1 TOC-LCSH/CSN Construction

2.1.3 Left Side View

Sampling tubing

Front door

IC reagent Dilution water bottle Drain tubing

Drain pot Hydrochloric acid

Figure 2.3 Left Side View

TOC-LCSH/CSN 9

2.1 TOC-LCSH/CSN Construction

2.1.4 Rear View

AC cord socket

Carrier gas inlet

Dip switches for various settings

Standalone/PC control switch

Printer connection socket

PC connection socket

ASI-L/OCT-L connection socket Extension board slot cover

Figure 2.4 Rear View

10 TOC-LCSH/CSN

2.1 TOC-LCSH/CSN Construction

2.1.5 Front Internal View (Left Diagonal)

Sparge gas flow meter

Sliding IC sample injector

(H Type only)

IC reaction vessel

(H Type only)

3-way tap

8-port valve

5 mL syringe

Cooling tubing unit

Solenoid valve (for IC drain discharge)

(H Type only)

(option)

Sparge gas flow adjustment

knob (option)

Carrier gas pressure

adjustment knob (option)

Halogen scrubber

Humidifier

Pump (for supplying IC reagent)

(H Type only)

Figure 2.5 Front Internal View (Left Diagonal)

TOC-LCSH/CSN 11

2.1 TOC-LCSH/CSN Construction

2.1.6 Top Internal View (Left Diagonal)

Motor for driving the slider

Sample injector block

Cooling fan for the sample injector block

Figure 2.6 Top Internal View (Left Diagonal)

2.1.7 Keyboard Layout

Function keys

MEAS key

ASI key

CAL key

METHOD key

SAVE METHOD

key

Power switch

START key

PRINT key

ABC key

NEXT key

BS key

Arrow keys

CE key

STOP key ENTER key Numeric keypad

Figure 2.7 Keyboard Layout

12 TOC-LCSH/CSN

2.1 TOC-LCSH/CSN Construction

Entering English Characters

Pressing the [ABC] key allows entry of English characters and symbols using the numeric  keys. Each English letter is assigned to a key, and each time the key is pressed, the letter  changes as indicated in the following table.

Numeric Key Corresponding English Letter / Symbol

1 . # /

2 A B C

3 D E F

4 G H I

5 J K L

6 M N O

7 P Q R S

8 T U V

9 W X Y Z

0 Moves cursor left

. Moves cursor right

TIP » Moving the Cursor

After displaying the character you wish to enter, pressing another key automatically  moves the cursor to the right, allowing entry of the next character. In addition, if the  next character to be entered is assigned to the same numeric key, first press "." to move  the cursor right.

Example: Entering "ABC"

1. Press the [ABC] key.

The [ABC] key illuminates.

2. Press "2" once.

The letter A is entered.

3. Press ".".

The cursor moves one space to the right.

4. Press "2" twice.

The letter B is entered.

5. Press ".".

The cursor move one space to the right.

6. Press "2" three times.

The letter C is entered.

7. Press the [Enter] key to confirm the entry.

TIP » Deleting Characters

To delete one character, press the [BS] key (the character at which the cursor is  positioned is deleted.) To delete all of the entered characters, press the [CE] key.

TIP » Inserting Characters

Move the cursor using the "0" key or ".", as appropriate, and then enter the

appropriate character.

TOC-LCSH/CSN 13

2.1 TOC-LCSH/CSN Construction

TIP » Adjusting the Screen Contrast

While keeping the ABC key in the "Initial Display" ON, adjust the contrast with the ↑ ↓ keys.

2.1.8 Flow Diagrams

Diluent Acid

Sample

Backflow  Prevention  Valve

Sliding TC  Sample  Injector

IC Reaction

Vessel Backflow  Prevention Trap

Solenoid

Valve 3

(To Drain)

IC Reagent

IC Reagent

Supply Pump

Phosphoric

Acid

CO2 Absorber

Light Source

NDIR Tandem Cell

(Sample Cell)

Halogen Scrubber

5 mL Syringe

Pressure Sensor PS3

Proportional Valve 2

Pure Water Trap

TC Combustion

Tube

Dehumidifier

(Electronic Cooler)

Drain Pot

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure  Sensor PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier

Cooling Tubing

Figure 2.8 Flow Diagram (TOC-LCSH)

Figure 2.9 Flow Diagram (TOC-LCSH + TNM-L + External Sparge Kit)

14 TOC-LCSH/CSN

2.1 TOC-LCSH/CSN Construction

Diluent Acid

Sample

IC Reagent

Supply Pump

Phosphoric

Acid

CO2 Absorber

Backflow  Prevention  Valve

Backflow

Prevention Trap

Solenoid

Valve 3

(To Drain)

Sliding TC

Sample Injector

IC Reaction

Vessel

IC Reagent

CO2 Absorber

Light Source

NDIR Tandem

Cell (Sample

Halogen Scrubber

Cell)

5 mL Syringe

Pressure Sensor PS3 Proportional Valve 2

Oxygen Permeation  Tubing

N2 Carrier Gas Kit

TC Combustion  Tube

Cooling Tubing

Pure Water Trap

Dehumidifier

(Electronic Cooler)

Drain Pot

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure  Sensor PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier

Figure 2.10 Flow Diagram (TOC-LCSH + N2 Carrier Gas Kit) IC Reaction

Diluent

Vessel Backflow

IC Reagent

4

5

Prevention Trap

Supply Pump

Acid

3 P

Sample or ASI

2

1

6

8 7

Solenoid Valve 4

Sliding TC

Solenoid

Valve 3

(To Drain)

IC Reagent

Phosphoric

Light Source

Solenoid Valve 2

CO Acid 2

(For Sparging)

CO2 Absorber

Absorber

Backflow  Prevention

Sample  Injector

Valve

Halogen Scrubber

NDIR Tandem  Cell (Sample  Cell)

5 mL Syringe

Pressure Sensor PS3

Proportional Valve 2

TC Combustion  Tube

Dehumidifier

(Electronic Cooler)

Pure Water Trap

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure Sensor  PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier Cooling Tubing

Drain Pot

Figure 2.11 Flow Diagram (TOC-LCSH + POC Kit)

TOC-LCSH/CSN 15

2.1 TOC-LCSH/CSN Construction

Manual IC

Gas Injector

Diluent

Acid

Manual TC

Gas Injector

Light Source

IC Reaction

Backflow Vessel Prevention Trap

Solenoid

Valve 3

(To Drain)

IC Reagent  Supply Pump

Sample

Backflow  Prevention  Valve

Sliding TC  Sample

Injector

IC Reagent

Phosphoric  Acid

CO2 Absorber

NDIR Tandem  Cell (Sample

Halogen Scrubber

Cell)

5 mL Syringe

Pressure Sensor PS3

Proportional Valve 2

Pure Water Trap

TC Combustion

Tube

Dehumidifier

(Electronic Cooler)

Drain Pot

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure  Sensor PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier

Cooling Tubing

Figure 2.12 Flow Diagram (TOC-LCSH + Manual Injection Kit)

Diluent

Acid

Sample

Light Source

Solenoid Valve 2 (For Sparging)

Backflow  Prevention  Valve

Sliding TC  Sample

Injector

Halogen Scrubber

CO2 Absorber

NDIR Tandem  Cell (Sample  Cell)

5 mL Syringe

Pressure Sensor PS3

Proportional Valve 2

Pure Water Trap

TC Combustion

Tube

Dehumidifier

(Electronic Cooler)

Drain Pot

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure  Sensor PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier Cooling Tubing

Figure 2.13 Flow Diagram (TOC-LCSN)

16 TOC-LCSH/CSN

2.1 TOC-LCSH/CSN Construction

Figure 2.14 Flow Diagram (TOC-LCSN + TNM-L + External Sparge Kit)

Solenoid

Diluent Acid

Sample

Solenoid Valve 2 (For Sparging)

CO2

Absorber

Valve 4

Backflow  Prevention  Valve

Sliding TC  Sample

Injector

Halogen Scrubber

CO2 Absorber

Light Source

NDIR Tandem

Cell (Sample Cell)

5 mL Syringe

Pressure Sensor PS3

Proportional Valve 2

Pure Water Trap

TC Combustion

Tube

Dehumidifier

(Electronic Cooler)

Membrane Filter Detector

Solenoid Valve 1

Carrier Gas Inlet

Pressure  Sensor PS1

Proportional  Valve 1

Pressure  Sensor PS2

Humidifier Cooling Tubing

Drain Pot

Figure 2.15 Flow Diagram (TOC-LCSN + POC Kit)

TOC-LCSH/CSN 17

2.1 TOC-LCSH/CSN Construction

Figure 2.16 Flow Diagram (TOC-LCSN + TNM-L + Carrier Gas Purification Kit)

18 TOC-LCSH/CSN

2.2 Autosampler (Option)

2.2 Autosampler (Option)

2.2.1 Front View (Left Diagonal)

Sparging tubing (option)

Sampling tubing

Door for needle replacement

Sample cover

Power switch

Figure 2.17 Front View (Left Diagonal)

2.2.2 Rear View (Left Diagonal)

Rinse bottle

Stirrer rotation speed

adjustment dial (option)

Drain tubing Tubing to supply rinse water Connection cable

Figure 2.18 Rear View (Left Diagonal)

TOC-LCSH/CSN 19

2.3 8-Port Sampler OCT-L (Option)

2.3 8-Port Sampler OCT-L (Option)

2.3.1 Front View (Left Diagonal)

8-port valve

Figure 2.19 Front View (Left Diagonal)

2.3.2 Rear View (Right Diagonal)

Connection cable

Connection cable

(for linking two OCT-L units)

Figure 2.20 Rear View (Right Diagonal)

20 TOC-LCSH/CSN

2.4 TN Unit (Option)

2.4 TN Unit (Option)

2.4.1 Front View (Right Diagonal)

TNM-L

TNM-L maintenance door

Figure 2.21 Front View (Right Diagonal)

2.4.2 View Inside the Door on the Right Side

Flow rate adjustment knob

for the ozone air source

Figure 2.22 View Inside the Door on the Right Side

TOC-LCSH/CSN 21

2.4 TN Unit (Option)

2.4.3 Rear View

Ozone air source inlet

Figure 2.23 Rear View

22 TOC-LCSH/CSN

3

3Operation

This chapter describes pre-analysis preparations, parameter entry procedures and  operation procedures for calibration curve generation and sample analysis.

3.1 Analysis Preparation

This section describes catalyst filling, combustion tube connection, syringe  installation and other procedures used for actual sample analysis.

3.2 Analysis

This section describes procedures from instrument start-up to calibration curve  generation. These procedures include sample analysis, method development,  halting of analysis and instrument shutdown.

3.3 Analysis with the Autosampler

This section describes the procedure for performing analysis using the  autosampler.

3.4 Analysis with the 8-Port Sampler

This section describes the procedure for performing analysis using the 8-port  sampler.

3.5 Data Review

Procedures for reviewing the calibration curve list, method list and data reports  are described here.

3.6 Outputting Data to a USB Memory Device

This section describes how to output data to a USB memory device.

3.7 Applications Operations (H Type Only)

The procedures for performing manual injection analysis and gas sample  analysis are described in this section.

33

3.1 Analysis Preparation

3.1 Analysis Preparation

3.1.1 Catalyst Filling

There are three types of catalyst: TOC standard catalyst, TOC high sensitivity catalyst (H  Type only), and TOC/TN catalyst (option). For TOC analysis, fill the combustion tube with  catalyst as described in Section 3.1.1.1 "Catalyst for TOC Analysis" on page 25. For TN  analysis, refer to Section 3.1.1.2 "TOC/TN Catalyst (Option)" on page 27.

CAUTION

Exercise caution when handling the combustion tube to prevent breakage as it is

constructed from quartz glass.

Note: When filling the combustion tube, verify that the catalyst does not become contaminated. Contamination will cause high blank values and abnormal analysis

values.

Using New Catalyst

Broad blank peaks are associated with the initial use of new catalyst. These blank peaks  become progressively smaller with repeated injections, and the peak size becomes stable.  This tendency is more obvious with high sensitivity TC analysis. If calibration or analysis  is performed while the blank peaks are large, the measurements are altered with the  changing magnitude of the blank peaks, resulting in errors in the analysis values.

TC Analysis Using New Catalyst

Catalyst regeneration should be conducted 1-2 times for TC analysis using new catalyst.  This reduces the size of the blank peaks. Use continuous injections of pure water until the  blank peak area becomes small and stable. Refer to Section 3.2.3.1 "Single-mode Analysis"  on page 114 step " (2) Continuous Measurement in Same Mode". The blank peak

magnitude depends on the quality of the water (TC impurity concentration) and the TC  range used for analysis. For analysis of concentrations 100mg/L or greater, a blank peak of  1mg/L will have little influence on the results. However, analysis of concentrations around  10mg/L must have stable blank peaks of 0.5mg/L or less.

In addition, measure the NPOC of the sample using the calibration curve created in the  previous section when using TOC high sensitivity catalyst. Measure the sample using the  same method used to measure standard solution, as described in (2) in the previous section.  Be aware that when sparging is performed using the optional external sparging kit and  repeated analysis of the same sample results in gradually decreasing values, the sparging  time may be too short. In this case, extend the sparging time. The required sparging time  depends on the shape of the sample container, sample volume, sparge gas volume, and  amount of IC in the sample. Determine a suitable sparging time according to each of these  conditions.

Perform the procedure described in Section 3.1.19.2 "TC Blank Check Analysis" on page  81.

24 TOC-LCSH/CSN

3.1 Analysis Preparation

TN Analysis Using New Catalyst

Blank peak magnitude poses no practical problems for TN analysis using new catalyst. The  blank peaks are fairly small and stable. Perform analysis with distilled or deionized water to  confirm the blank peak size.

3.1.1.1 Catalyst for TOC Analysis

Either standard TOC catalyst or optional TOC high sensitivity catalyst can be used for TOC  analysis. The furnace temperature is set to 680°C for TOC analysis.

⮴TOC Standard Catalyst

The commonly used standard TOC catalyst is appropriate for both the TC-IC method and  NPOC method of analysis.

Note: Analyze TOC high precision samples of about 1mg/L using the TOC high sensitivity catalyst because smaller blank peaks are generated.

Filling Combustion Tube with TOC Standard Catalyst

1. Insert two sheets of platinum mesh into the combustion tube.

2. Layer about 5mm of quartz wool onto the platinum mesh using the catalyst filling rod. 3. Fill the TOC standard catalyst to a height of about 100mm.

Combustion Tube

Standard TOC Catalyst

100mm

Quartz Wool

Platinum Mesh

Figure 3.1 Combustion Tube with Standard TOC Catalyst

TOC-LCSH/CSN 25

3.1 Analysis Preparation

CAUTION

Take care not to inhale any quartz wool when filling the combustion tube.

Inhaling quartz wool may cause health impairment. Handle with care. For

details, see Section 5.10 "Safety Data Sheets" on page 381

⮴TOC High Sensitivity Catalyst (H Type only)

Use the TOC high sensitivity catalyst is for low level TC analysis. The blank values are  controlled to negligible levels, and analysis is performed primarily for trace TOC in pure  water (1mg/L or less).

Samples containing IC in excess of 5mg/L The detection rate for IC decreases with increasing concentration, therefore, when IC is contained in the sample in excess of 5mg/L, use the NPOC method instead of the TC-IC method.

Filling Combustion Tube with TOC High Sensitivity Catalyst

Two cylindrical containers of catalyst are required for a single load. The procedure is as  follows.

1. Place two sheets of platinum mesh in the combustion tube.

2. Fit together the combustion tube opening and the opened end of the cylindrical

catalyst container.

3. Use the catalyst filling rod to push catalyst into the combustion tube from the other  opened end of the cylindrical catalyst container.

4. Transfer the contents of the two cylindrical catalyst containers to achieve a catalyst  layer of 110 mm in the combustion tube.

A TC high sensitivity catalyst kit, including a combustion tube filled with TOC high sensitivity catalyst in advance, is also available.

26 TOC-LCSH/CSN

3.1 Analysis Preparation

Combustion Tube

TOC High Sensitivity

Catalyst

110 mm

Platinum Mesh

Figure 3.2 Filling of TOC Combustion Tube with High Sensitivity Catalyst

CAUTION

Wear protective glasses and gloves when handling high sensitivity catalyst.

Inhaling it may cause health impairment.

3.1.1.2 TOC/TN Catalyst (Option)

TOC/TN catalyst is used for TN analysis and TN/TOC simultaneous analysis. TOC/TN  catalyst is not applicable for TC-IC analysis because the detection rate for the IC

component in excess of 10mg/L is low. Use the NPOC method for TOC analysis in these  situations. The furnace temperature is set to 720°C.

Filling Combustion Tube with Catalyst for TN Analysis

1. Insert two sheets of platinum mesh into the combustion tube.

2. Layer about 5mm of quartz wool onto the platinum mesh using the catalyst filling rod. 3. Introduce the standard TOC catalyst into the combustion tube to 140mm from the top  end of the combustion tube.

4. Introduce 0.2 (± 0.1) g of ceramic fiber to cover the catalyst.

TOC-LCSH/CSN 27

3.1 Analysis Preparation

5. Lightly press down on the ceramic fiber using the catalyst filling rod to attain a  thickness of 10 mm.

Note: • The thickness of the ceramic fiber must be uniform, covering the entire  surface of the catalyst, or a poor TN detection rate may result.

• The filling heights of the catalyst and ceramic fiber influence the TN

detection rate, so it is important that the dimensions are close to those

indicated below.

Catalyst: 140mm from the top end of the combustion tube

Ceramic fiber: 10mm

Combustion Tube

Ceramic Fiber

Standard TOC Catalyst Quartz Wool

10mm

140mm

Platinum Mesh

Figure 3.3 Filling Combustion Tube with Catalyst for TOC/TN Analysis

CAUTION

Take care not to inhale any quartz wool or ceramic fiber when filling the  combustion tube.

Inhaling quartz wool may cause health impairment. Handle with care. For  details, see Section 5.10 "Safety Data Sheets" on page 381.

28 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.2 Installing the Cooling Tube Unit (H Type Only)

Install the cooling tube unit to the outlet of the combustion tube.

CAUTION

• The cooling tube unit is made of hard glass. Handle the unit with care to prevent breakage.

• Tighten the hexagon cap nut while holding the PTFE coupling in place with your finger when connecting the outlet tube to the cooling tube. The cooling tube with

break if the hexagon nut is tightened without holding the PTFE coupling in place.

1. Open the front door on the instrument and connect the PTFE tube attached to the pure  water trap to the outlet of the cooling tube unit (the smaller PTFE coupling). At this  point, carefully tighten the hexagon cap nut while holding the PTFE coupling in place  with your finger.

2. Loosen and remove the two fastening screws on the cooling tube unit located on the  lower section of the electric furnace and insert the screws into each of the two holes  on the upper section of the cooling tube unit and tighten them.

To the pure water trap

Cooling tube unit outlet

(PTFE coupling)

Fastening

screws

Cooling tubing unit

Figure 3.4 Installing the Cooling Tube Unit

TOC-LCSH/CSN 29

3.1 Analysis Preparation

3.1.3 Connecting the Combustion Tube

The procedure for installing the combustion tube differs depending on whether the

instrument is the H Type or the N Type.

CAUTION

• Due to the potential for burns, maintenance of the combustion tube and catalyst should be performed after the electric furnace temperature has cooled to room

temperature.

• When connecting the combustion tube, handle it carefully to avoid breaking it. The combustion tube is constructed of quartz glass, and injury is possible.

• Do not use a wrench or other tool. Injury could result if the combustion tube breaks.

Installation for the H Type

1. Open the door on top of the instrument, remove the slider retaining screws (knurled  screws) from the sample injector, and pull the slider towards you and out of the

sample injector.

2. Loosen the hexagon cap nut that joins the carrier gas introduction tubing (PTFE  tubing) to the right side of the sample injector and then remove the tubing.

3. Loosen the two knurled screws on the sample injector and remove the sample injector  in the upward direction.

Slider retaining

screw

Slider

Sample injector retaining screws (two)

Figure 3.5 Sample Injector

Drain tubing

Carrier gas

introduction tubing Hexagon cap nut Sample injector

4. Securely insert the top end of the combustion tube filled with catalyst into the hole on  the bottom of the sample injector.

30 TOC-LCSH/CSN

3.1 Analysis Preparation

5. Fit the combustion tube into the hole in the center of the electric furnace and then  insert the bottom end of the combustion tube into the PTFE coupling of the cooling  tubing inlet. At this point, securely connect the drain tubing, which protrudes out of  the sample injector, to the Y-shaped fitting.

Sample injector

Drain tubing

Combustion

tube

Y-shaped fitting

Figure 3.6 Installing the Combustion Tube

6. Secure the sample injector to the instrument body using the knurled screws.

7. Reattach the slider and carrier gas introduction tubing to their original positions.

Sample injector

Drain tubing

Sample injector retaining screws

Figure 3.7 Securing the Sample Injector to the Instrument Body

TOC-LCSH/CSN 31

3.1 Analysis Preparation

Note: Do not allow organic substances, such as grease or oil, to contaminate any parts exposed to gas including the inner surfaces on the tip of the combustion tube as

well as joints. This is especially important for high sensitivity measurement.

8. Connect the bottom of the combustion tube to the PTFE coupling of the cooling  tubing inlet. Carefully tighten the hexagon cap nut while holding the PTFE coupling  in place with your finger.

Combustion tube

Cooling tubing unit

PTFE coupling

(cooling tubing inlet)

Figure 3.8 Connecting the Combustion Tube and Cooling Tubing

Bottom end of the

combustion tube

Hexagon cap nut

PTFE coupling

Cooling tubing inlet

Figure 3.9 Connecting the Bottom End of the Combustion Tube to the PTFE Coupling

CAUTION

The quartz glass combustion tube will break if the hexagon cap nut is

overtightened.

Tightening the PTFE coupling without holding it in place will break the cooling

tubing because it is made of hard glass.

32 TOC-LCSH/CSN

3.1 Analysis Preparation

9. When using catalyst for TOC measurement and TOC/TN catalyst, use the provided  quartz wool (0.6 g and more) to fill the gap between the bottom of the electric furnace  and the combustion tube as well as the two notches in the insulator on top of the

electric furnace.

CAUTION

• When using TOC high sensitivity catalyst, do not fill the gap at the bottom of the electric furnace with quartz wool.

• Take care not to inhale any quartz wool when filling the gap and the notches.

• Inhaling quartz wool may cause health impairment. Handle with care. For details, see Section 5.10 "Safety Data Sheets" on page 381.

Combustion  tube

(Standard  catalyst for  TOC,

Catalyst for  TC/TN)

Combustion tube

(High sensitivity

catalyst)

Electric

furnace

Quartz wool

(0.6 g and more)

Figure 3.10 Filling With Quartz Wool (Bottom of the Electric Furnace) Fill with quartz wool InsulatorFigure 3.11 Filling With Quartz Wool (Top of the Electric Furnace)

Electric  furnace

TOC-LCSH/CSN 33

3.1 Analysis Preparation

Installation for the N Type

1. Open the door on top of the instrument, remove the slider retaining screws (knurled  screws) from the sample injector, and pull the slider towards you and out of the

sample injector.

2. Loosen the hexagon cap nut that joins the carrier gas introduction tubing (PTFE  tubing) to the right side of the sample injector and then remove the tubing.

3. Loosen the two knurled screws on the sample injector and remove the sample injector  in the upward direction.

Slider retaining

screw

Slider

Sample injector retaining screws (two)

Figure 3.12 Sample Injector

Drain tubing

Carrier gas

introduction tubing Hexagon cap nut Sample injector

4. Securely insert the top end of the combustion tube filled with catalyst into the hole on  the bottom of the sample injector.

34 TOC-LCSH/CSN

3.1 Analysis Preparation

5. Fit the combustion tube into the hole in the center of the electric furnace and then  insert the bottom end of the combustion tube into the PTFE coupling of the cooling  tubing inlet. At this point, securely connect the drain tubing, which protrudes out of  the sample injector, to the Y-shaped fitting.

Sample injector

Drain tubing

Combustion

tube

Y-shaped fitting

Figure 3.13 Installing the Combustion Tube

6. Secure the sample injector to the instrument body using the knurled screws.

7. Reattach the slider and carrier gas introduction tubing to their original positions.

Sample injector

Drain tubing

Sample injector retaining screws

Figure 3.14 Securing the Sample Injector to the Instrument Body

TOC-LCSH/CSN 35

3.1 Analysis Preparation

Note: Do not allow organic substances, such as grease or oil, to contaminate any parts exposed to gas including the inner surfaces on the tip of the combustion tube as

well as joints. This is especially important for high sensitivity measurement.

CAUTION

The quartz glass combustion tube will break if the compression fitting is over

tightened.

8. Connect the bottom of the combustion tube to the PTFE coupling of the cooling  tubing inlet. Carefully tighten the hexagon cap nut while holding the PTFE coupling  in place with your finger.

Cooling tubing PTFE coupling (Connects to the

bottom end of the combustion tube.)

Figure 3.15 Connecting the Combustion Tube and Cooling Tubing

36 TOC-LCSH/CSN

3.1 Analysis Preparation

9. In order to prevent heat from escaping from the furnace, use the provided quartz wool  (0.6 g and more) to fill the gap between the bottom of the electric furnace and the

combustion tube as well as the two notches in the insulator on top of the electric

furnace.

CAUTION

Take care not to inhale any quartz wool when filling the combustion tube.

Inhaling quartz wool may cause health impairment. Handle with care. For

details, see Section 5.10 "Safety Data Sheets" on page 381.

Combustion tube

Electric furnace

Quartz wool

(0.6 g and more)

Figure 3.16 Filling With Quartz Wool (Bottom of the Electric Furnace)

Fill with quartz wool Insulator

Figure 3.17 Filling With Quartz Wool (Top of the Electric Furnace)

TOC-LCSH/CSN 37

3.1 Analysis Preparation

CAUTION

When installing or moving the instrument, insulation material powder may drop

under the electric furnace. Since insulation material contains ceramic fibers,

follow these precautions:

• When cleaning dust under the electric furnace, wipe with dampened cloths to

prevent dust or powder from being dispersed.

• Ceramic fibers are contained in the material used inside the electric furnace.

Never dismantle the electric furnace.

Refractory ceramic fibers (RCF) are fire-resistant inorganic fibers mainly consisting of alumina (Al2O3)  and silica (SiO2). This material is chemically stable and can be used as insulation material in a high  temperature range exceeding 1000 °C.

With this instrument, the dispersibility of fibrous mineral under normal conditions of use is measured and  is proven to be sufficiently low with respect to the maximum permissible dust concentration level.

38 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.4 Connecting the Halogen Scrubber

Connection Procedure

1. Remove the film at both ends of the new halogen scrubber and connect it to the Viton  connectors.

2. Secure the halogen scrubber with two white plastic clamps.

Replace Scrubber When

Discoloration Band Reaches

Dehumidifier

Gas Flow Direction

To Sample

Cell Inlet

Halogen Reactant-filled Cylinder

this Point

Figure 3.18 Halogen Scrubber

The baseline may fluctuate immediately after connection. Allow the instrument to run for a  short time until the baseline stabilizes. Stabilization normally occurs within one hour.

Halogen scrubber

(take note of the

arrow direction)

Figure 3.19 Halogen Scrubber Connection

TOC-LCSH/CSN 39

3.1 Analysis Preparation

3.1.5 Installing the Syringe

This section describes how to install the syringe into the instrument.

Each port of the 8-port valve for the syringe pump type sample injector is connected to  certain streams depending on the state of the instrument. Installing and removing the  syringe without following the correct procedure may cause acid and other substances to be  discharged from the syringe connection port. Always follow the procedure described  below.

CAUTION

• If the syringe is installed using a procedure other than that described below, there is a danger that acid, or other solutions may leak from the syringe connection

port. Ensure that the syringe is installed according to the procedure below.

• To avoid injury, keep hands away from the sample injector during operation of the syringe pump sample injector.

• The syringe barrel is made of glass, handle it carefully to avoid breakage.

Syringe Replacement Procedure

1. In the "Initial Display", press the F5 [Maintenance] key.

The "Maintenance" screen is displayed.

2. Using the ↓ key, move the cursor to the "Replace Syringe" item.

3. Press the [NEXT] key.

When the preparations for syringe replacement are completed, the message "Syringe  exchange. Exchange syringe now." is displayed.

40 TOC-LCSH/CSN

3.1 Analysis Preparation

4. Attach the syringe to the syringe connection port (downward facing port) of the 8-port  valve and then connect the sparge gas line. At this point, the plunger retaining screw  is not completely tightened. Tighten the retaining screw so that the plunger can still be  slightly moved forwards and backwards as well as left and right.

Note: Do not use tools when replacing the syringe, hand tighten only.

Over-tightening can deform the resin components of the 8-port

valve, resulting in leaks. As a guide, tighten the screw until it stops

snugly and make one quarter turn further.

The plunger is secured to the plunger holder with the plunger

retaining screw through a spring washer. Be careful not to lose the

spring washer.

5. Press the F4 [Move test Up & Down] key.

The plunger holder will rise.

6. Once the syringe reaches the top position, hold the plunger vertical and tighten the  plunger retaining screw to secure it to the plunger holder.

7. Connect the sparge gas line to the syringe using the bushing and spacer.

8. Press the F6 [End] key.

The operation finishes, and the display returns to the "Maintenance" screen.

9. Press the F1 [Return] key.

The display returns to the "Initial Display".

8-port valve

Syringe

Plunger Plunger retaining screw

Plunger holder

Figure 3.20 Replacing the Syringe

TOC-LCSH/CSN 41

3.1 Analysis Preparation

3.1.6 Water Supply to the Dehumidifier Drain Vessel

The drain container must be filled with water to prevent carrier gas release from the drain  tubing.

Procedure

1. Remove the cooler drain vessel from the mounting clip and then remove the rubber  cap from the vessel.

2. Using a wash bottle, pour pure water (deionized water is acceptable) into the cooler  drain vessel up to the level of the drain discharge tubing on the side of the drain

vessel.

3. Reattach the rubber cap onto the cooler drain vessel and return the vessel to its  position in the mounting clip.

4. Verify that the cooler drain tubing is inserted straightly in the drain vessel without  being bent.

CAUTION

Do not bend drain tubing, as this may prevent drain discharge.

Rubber cap

Cooler drain vessel

Drain discharge tube

Figure 3.21 Filling the Dehumidifier Drain Vessel with Water

42 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.7 Filling the Humidifier with Water

It is necessary to humidify the carrier gas for high sensitivity analysis, to ensure high  analysis precision. Water is supplied to the humidifier using the following procedure.

Procedure

1. Remove the cap from the water supply port at the top of the humidifier.

2. Introduce water through the water supply port to the "Hi" marking on the side of the  humidifier.

Note: Use pure water containing as little TC as possible for high

sensitivity analysis.

Water supply port

Humidifier

Figure 3.22 Filling the Humidifier with Water

TOC-LCSH/CSN 43

3.1 Analysis Preparation

3.1.8 Preparation of Dilution Water

This section describes how to prepare the dilution water required to automatically dilute the  sample inside the instrument.

Filling Procedure

1. Wash the provided dilution water bottle (2 L capacity) with laboratory detergent, rinse  well with tap water and then a number of times with pure water.

2. Fill the dilution water bottle with pure water and place it to the left of the instrument. 3. Pass the dilution water intake tubing, labeled DILUTION, through the hole in the cap.  4. Attach a wiring band about 250 mm from the end of the dilution water intake tubing  to prevent the tubing from separating from the cap.

5. Place the cap on the dilution water bottle.

Note: Verify that the tip of the intake tubing nearly reaches the bottom of

the container.

Dilution water

intake tubing

Dilution water bottle

Figure 3.23 Preparing Dilution Water

Note: Pure Water

• Use the purest water that can be obtained.

• Refer to Section 3.1.18.1 "Preparation of Standard Solutions" on page 75

for details on the degree of purity of "pure water".

44 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.9 Acid Preparation

This section describes the preparation of acid used for IC removal in NPOC measurement  and IC measurement with the N Type instrument.

Although the concentration of hydrochloric acid is normally 1 mol/L as described below,  adjust the concentration according to the properties of the sample.

Preparation Procedure

1. Dilute commercially available special grade concentrated hydrochloric acid

(12 mol/L) with 12 parts pure water to prepare 1 mol/L hydrochloric acid. A

concentration accuracy of approximately ±2 % is acceptable.

The pure water used as the diluent must have a TOC concentration no greater than

0.5 mg/L. Since the TOC value of the diluent pure water can have a large effect on the  analysis result with high sensitivity analysis (low TOC concentrations), it is necessary  to use pure water having a very low TOC concentration.

CAUTION

Hydrochloric acid is a hazardous chemical. Use precautions to avoid contact

when handling this substance. For details, refer to Section 5.10 "Safety Data

Sheets" on page 381.

2. Pour the prepared acid into the 250 mL acid container provided standard with the  instrument, and secure the acid container in the mounting clip on the left side of the  instrument.

3. Pass the acid intake tubing through the hole in the cap.

4. Secure the tubing, using a wire clamp about 100 mm from the end of the tubing, to  prevent the tubing and cap from separating.

5. Place the cap on the container.

Note: Verify that the tip of the acid intake tubing nearly reaches the

bottom of the container.

Acid intake tubing

Acid container (250 mL capacity)

Figure 3.24 Acid Preparation

TOC-LCSH/CSN 45

3.1 Analysis Preparation

3.1.10 Installation of the CO2 Absorber

A CO2 absorber is required to remove CO2 from carrier gas because carrier gas is used to  purge the optical system of the detector. Install the CO2 absorber by following the

procedure described below.

Procedure

1. Cut the tips of the rigid tubes (2) in the lid of the CO2 absorber with a knife.

2. Connect the flexible tubing that exits the rear panel of the instrument near the top as  described below.

Tubing labeled "L": Connect to tube labeled "L" on CO2 absorber.

Tubing labeled "S": Connect to tube labeled "S" on CO2 absorber.

Note: The flexible tubing should not be bent.

3. Place the CO2 absorber in the holder on the upper rear of the instrument.

Connect to

the absorber

CO2 absorber

Figure 3.25 Installation of the CO2 Absorber

3.1.11 Printer Preparation

Printers having the following control codes can be connected to the TOC-L.

ESC/P-Raster (some models)

ESC/P-9

ESC/P-24 (ESC/P-J84 or ESC/P Super (ESC/P emulations))

PCL (PCL3, PCL5, PCL5e)

Printer Connection Procedure

1. Connect the printer by plugging the USB cable into the printer connection terminal on  the rear of the TOC-L.

2. See Section 3.1.17.5 "Printer Setup" on page 62 and set up the printer.

46 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.12 Installing the Suspended Solids Kit

To analyze samples containing suspended solids, the sample flow line diameter can be  increased from 0.5mm to 0.8mm by installing the optional Suspended Solids Kit.

This section describes the installation of the Suspended Solids Kit.

There are two kits available:

1) For TOC-L (TOC-L main unit only)

2) For TOC-L+ASI-L (used with TOC-L and ASI-L)

CAUTION

Switch OFF the TOC-L before attempting to install the Suspended Solids kit.

⮴Replacing the Flow Line Parts and Slider

This procedure is used for both kits.

Procedure

1. Remove the bushing from port 7 of the 8-port syringe pump valve, and disconnect the  TC injection tubing. For the H-type, remove the bushing from port 5, and disconnect  the IC injection tubing.

2. Remove the slider mounting screws from the TC and IC injection blocks, and detach  the sliders.

3. Remove the slider bushings and replace the sample injection tubing with the 0.8 mm  diameter sample injection tubing provided in the kit.

4. Re-install the sliders on the sample injection blocks.

5. Connect the TC injection tubing and the IC injection tubing to port 7 and port 5,  respectively.

Note: During analysis, verify that the sample flow is normal. Refer to

Section 4.6.2.4 "Remedial Actions for Poor Repeatability" on page

277.

If the sample flow is not normal, the tip of the PTFE tube may be deformed

or scratched. Use a knife to cut off a portion of the tube.

⮴Replacing Sample Aspiration Tubing

For TOC-L (main unit only)

1. Remove the bushing from port 1 of the 8-port syringe pump valve, and replace the  sample aspiration tubing with the 0.8mm diameter tubing provided in the kit.

2. Connect the new sample aspiration tubing to the 8-port valve in the original position.

TOC-LCSH/CSN 47

3.1 Analysis Preparation

For TOC-L+ASI-L (TOC-L and ASI-L)

Refer to Section 5.9.4 "Installing the ASI-L" on page 369, to remove the autosampler  sample needle and sample aspiration tubing. Replace these with the sample needle and  aspiration tubing provided in the Suspended Solids kit.

Precautions for Measuring Suspended Solids

When performing measurement of solid samples using the Suspended Solids

Measurement kit, follow the precautions below to ensure the best possible

conditions for measurement. Refer to Section 3.1.19.4 "Analysis of Samples

Containing Suspended Solids" on page 84.

1) Before measurement, use a homogenizer, etc. to make the sample as uniform as  possible. After homogenizing the sample, depending on the type of suspended

substance, precipitation and coagulation may occur after a period of time. Prior

to measurement, mix the sample again using a magnetic stirrer to ensure that

precipitated/coagulated material is well dispersed throughout the sample.

2) If the suspended material is present in large quantities, clumping of the material  and obstruction may occur inside the sample flow line even if the particles are of

smaller diameter that the sample flow line diameter. Take measures such as

sample dilution to prevent flow line obstruction,

3) The suspended solids measurement kit employs a 0.8 mm PTFE tubing for

sample injection. Diminished reproducibility may occur with small sample

injection volumes. Make injection volumes as large as possible.

4) When conducting IC measurement of suspended solid samples using the H type  instrument, obstruction of the IC measurement flow line drain tube may occur.

Sufficiently pulverize the suspended particles to allow free flow through the

drain tube.

48 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.13 IC Analysis Preparations

Certain preparations are necessary for the H Type instrument. These include installing and  replenishing the IC reagent in the IC reaction vessel. This section describes the procedures  for installing and replenishing the IC reagent.

CAUTION

The IC reagent contains phosphoric acid, use care while handling to prevent

contact or spills. For details refer to Section 5.10 "Safety Data Sheets" on page

381.

⮴Installing the IC Reagent

Procedure

1. Prepare the IC reagent by diluting 50mL of commercial 85% phosphoric acid (ACS  Reagent grade) to a final volume of 250mL with pure water.

Note: Since the specific gravity of 85% phosphoric acid is about 1.7, the

concentration of the phosphoric acid is about 25% (weight%) in the

prepared IC reagent. A final concentration accuracy of ±2% is

acceptable.

2. Pour the prepared IC reagent into the 250 mL container provided standard with the  instrument and secure the container with the mounting clip on the left side of the

instrument.

3. Pass the tip of the IC reagent tubing through the hole in the cap.

4. Secure the tubing, using a wire clamp at a point about 100mm from the end of the  tubing tip, to prevent the tubing and cap from separating.

5. Replace the cap on the container.

TOC-LCSH/CSN 49

3.1 Analysis Preparation

⮴Replenishing IC Reaction Vessel with IC Reagent

Use the following procedure to replenish the IC reagent.

Procedure

1. Verify that the lever of the 3-way valve at the bottom of the IC reaction vessel is  positioned to the left as shown in Figure 3.26 "Lever Position of IC Reaction Vessel 3- Way Valve".

Figure 3.26 Lever Position of IC Reaction Vessel 3-Way Valve

2. Press the F5 [Maintenance] key in the "Initial Display".

The "Maintenance" screen is displayed.

3. Place the cursor at the IC Reagent Automatic Regeneration (Auto IC regeneration)  item, and press the [NEXT] key.

The IC reagent will be delivered to the IC reagent vessel by the liquid delivery pump. 4. Repeat step 3 until the IC reagent reaches the level of the IC reagent drain tubing  (bottom of branch).

50 TOC-LCSH/CSN

3.1 Analysis Preparation

3.1.14 Preparation for POC Analysis

POC (purgeable organic carbon) analysis can be performed if the optional POC analysis kit  is installed on this instrument. This section describes preparations required for POC  analysis.

Installation of the CO2 Absorber for POC Analysis

1. Remove the CO2 absorber.

CO2 absorber for

POC analysis

Figure 3.27 CO2 Absorber Container for POC Analysis

2. Insert about 5mm of quartz wool into the bottom of the container.

3. Add commercially available anhydrous lithium hydroxide to the container on top of  the quartz wool until it reaches just below the screw threads of the container.

Note: About 6g of lithium hydroxide crystals is required for a single load.

4. Place about 5mm of quartz wool on top of the lithium hydroxide.

5. Attach the O-ring and nipple to the top of the container, and firmly tighten the cap. 6. Install the CO2 absorber for POC analysis in the original position.

TOC-LCSH/CSN 51

3.1 Analysis Preparation

CAUTION

• Lithium hydroxide is available in anhydrous and hydrated forms. Since the ability of the hydrated form to absorb carbon dioxide is slightly inferior to that of

the anhydrous form, use anhydrous lithium to fill the CO2 absorber container.

• When introducing the quartz wool and lithium hydroxide into the container, do not pack too tightly.

• The CO2 absorber container is constructed of glass. Handle carefully to prevent breakage.

CAUTION

Wear protective glasses, gloves, and mask when handling anhydrous lithium

hydroxide.

Anhydrous lithium hydroxide is a corrosive (strongly alkaline) powder.

Direct contact or inhaling it causes injury.

CAUTION

Inhaling lithium hydroxide may cause health impairment. Handle with care. For

details, check the contents of the safety data sheet (SDS) available from the

lithium hydroxide manufacturer.

CAUTION

Inhaling quartz wool may cause health impairment. Handle with care. For

details, see Section 5.10 "Safety Data Sheets" on page 381.

52 TOC-LCSH/CSN

3.1 Analysis Preparation

Cap

Nipple

O-Ring

Lithium Hydroxide

Quartz Wool

Figure 3.28 Preparing the CO2 Absorber for POC Analysis

TOC-LCSH/CSN 53

3.1 Analysis Preparation

3.1.15 TN Analysis Preparations (When Using the Optional TN Unit)

TN analysis is possible if the optional TN unit is installed with this instrument. The NOx absorber must also be connected to perform TN analysis.

3.1.15.1 NOx Absorber

The CO2 absorber (soda lime) is used as the NOx absorber. Connect it to the instrument  using the following procedure.

Procedure

1. Cut the tips of the rigid tubes (2) in the lid of the CO2 absorber with a knife.

2. Connect the Viton tubing to the 2 rigid tubes as described below.

Tubing labeled "L": Connect to tube labeled "L" on the CO2 absorber.

Tubing labeled "S": Connect to tube labeled "S" on the CO2 absorber.

3. Once connected, place the NOx absorber at the back right of the instrument.

Note: Make sure there are no kinks in the flexible tubing.

Connect to the absorber

NOx absorber

Figure 3.29 Connecting the NOx Absorber

54 TOC-LCSH/CSN